Michael Green  – Viamed Ltd

Armstrong Medical - Neoflow

First issue: 27th July 2015

Updated: 27th July 2015

Armstrong Medical - NeoFlow

The Armstrong medical NeoFlow ventilator circuit has been tested in conjunction with the Tom Thumb TT480 Infant resuscitator.

	
	Test at 5 lpm
	Test at 7 lpm
	Test at 10 lpm

	1 Out of Packet​​    
	6 cmH2O
	12 cmH2O
	27 cmH2O

	2 Max PEEP
	18 cmH2O
	28 cmH2O
	* 30 cmH2O

	3 Min PEEP
	0.6 cmH2O
	1 cmH2O
	6 cmH2O


 * Limited by precision safety valve.

1 Out of packet. The PEEP cap was pre-set to half a turn from maximum.

2 The maximum PEEP setting was achieved when the PEEP cap was rotated fully clockwise. 

3 The minimum PEEP setting was achieved at three full turns anti-clockwise from the maximum setting. The PEEP cap can be rotated further anti-clockwise without further effect.

The NeoFlow ventilator circuit has been tested with taper gauges. The 15mm Tom Thumb fitting tested within tolerance. The connection fitted very well when tested on a Tom Thumb.

The 15mm mask fitting was found to be tight but within the tolerance limit. The connection fitted very well when tested with a silicone mask.

