	CERTIFICATE OF CALIBRATION
Issued by:                   Vandagraph Ltd.
Date issued:               29 September 2015
Certificate No:              1443539065
Signed:                       ____________                              



Client: Vobster Marine Systems

Date of calibration. 
29 September 2015
Oxygen Monitor Type: VN202 Mk2



Serial number: 

00587

Procedure I.D.  

VM3COP40.56



Performance Status.  
PASS
The equipment detailed herein has been calibrated and now conforms to the manufacturer’s published specifications. The Q.A. of the above is in strict conformity to BS EN ISO 9001-2000 & ISO 13485-2003.  All measurements made are traceable to NAMAS.
	Equipment used to carry out calibration

	Fluke Graphical Multimeter.
	CE169

	Microcal  
	CE076

	Oxygen BOC Medical 99.5% License 07351
	

	Helium  BOC Medical 99.5% License 07351
	


	Test with Dry Air & 100% oxygen    Sensor S/n    
	565825

	100% Oxygen
	45.4mV
	
	

	Dry Air 20.9
	9.3mV
	
	

	Test with  Microcal
	DVM Reading
	Oxygen Target
	Oxygen Reading +/-1%

	Set Microcal to  +/- .05mV
	
	
	

	40mV
	40mV
	100
	100%

	30mV
	30mV
	75
	75%

	20mV
	20mV
	50
	49.9%

	10mV
	10mV
	25
	24.8%

	5mV
	4.98mV
	12.5
	12.3%

	1mV
	0.98mV
	2.5
	2.2%


Vandagraph Calibration Cert O2 Monitors 1443539065.doc                            25 June 2014

