Fault Finding on a Microstim DBIII

1. Remove the rear of the enclosure

2. Remove the 4 screws holding the pcn.

[image: image1.jpg]



3. Using long screws .Screw the test pcb on top of the circuit so the test prongs rest on the pcbUse Microstim test box

4. Using a DVM test for 9v on Batt in

5. Switch on 1HZ (i.e. continuous 1Hz) Test for approx 8.5v on D1 o/p 
If no output replace D1
6. With Oscilloscope set to   Vertical 1 v / div  Horizontal  100us /div  Ground to Batt -ve
7. Check Ref IN for 8.4 v approx Waveform should peak and then may drop slightly during pulse depending on condition of the battery returning to the original voltage at the start of each pulse. If it varies widely change the battery.

8. Test Ref out for 5v. Depending on age of battery it may drop slightly during pulse.

9. Check LED is Green and not going Amber or Red during the pulse.

10. Test Oisc Xtal for output
No output replace Xtal
11. Test Sig to LED for output
If no output change microprocessor
12. Test LED I/p for output
If no output change TR3 (Q1) 
13. Test Sig o/p for signal
If no output change microprocessor
14. Test o/p to TX for signal  Check TR1 & TR2 (Q2 Q3) or replace 
15. With Oscilloscope set to   XXXXXXXXXXXXXXXX  Test Audio feedback for signal

16. If no output change ZD3
17. Disconnect output load Audio should bleep.

18. Reconnect output load 
19. LED Checks

20. With Oscilloscope set to   XXXXXXXXXXXXXXXX  Ground to -ve battery

21. ZD1 waveform should be      v peak

22. ZD2 waveform should be      v peak
Microstim Fault finding.doc          Page 1                                      16. Oct. 2015

