Michael Green – Viamed Ltd
Flow sensor testing P/N 4310002
First issue: 24th August 2015
Updated: 24​​​​th August 2015

Comparison of flow sensors P/N 4310002
The following test data relates to flow Viamed flow sensors P/N 4310002.

It was noted that the difference between the top and bottom outputs did not vary significantly when compared to a Dräger Spirolog 8403735-10 flow sensor. The Viamed sensors did read lower overall in comparison to a sample batch of Dräger sensors that were tested. This difference would typically be 0.1 to 0.15 lower on the Viamed flow sensor.

	[image: image1.jpg]eaeins M ViIAmeD
i [T

8001005PCE





	[image: image2.jpg]





	Viamed flow sensor P/N 4310002

	LOT
	No.
	Test results @20°C

	
	
	Top
	Bottom
	Difference

	8001005PCB
	1
	3.257
	3.293
	-0.036

	8001005PCB
	2
	3.256
	3.261
	-0.005

	8001005PCB
	3
	3.294
	3.273
	0.021

	8001005PCB
	4
	3.297
	3.241
	0.056

	8001005PCB
	5
	3.314
	3.275
	0.039

	8001005PCB
	6
	3.284
	3.253
	0.031

	8001005PCB
	7
	3.291
	3.280
	0.011

	8001005PCB
	8
	3.227
	3.228
	-0.001

	8001005PCB
	9
	3.240
	3.268
	-0.028

	8001005PCB
	10
	3.289
	3.224
	0.065

	8001005PCB
	11
	3.251
	3.224
	0.027

	8001005PCB
	12
	3.262
	3.253
	0.009

	8001005PCB
	13
	3.279
	3.220
	0.059

	8001005PCB
	14
	3.212
	3.218
	-0.006


