	CERTIFICATE OF CALIBRATION
Issued by:                   Vandagraph Ltd.
Date issued:               25 November 2015
Certificate No:              1448450403
Signed:                       ____________                              



Client: Chesterfield Special Cylinders
Date of calibration. 
25 November 2015
Oxygen Monitor Type: VN202 mk2


Serial number: 

00760
Procedure I.D.  

VM3COP40.56



Performance Status.  
PASS
The equipment detailed herein has been calibrated and now conforms to the manufacturer’s published specifications. The Q.A. of the above is in strict conformity to BS EN ISO 9001-2000 & ISO 13485-2003.  All measurements made are traceable to NAMAS.
	Equipment used to carry out calibration

	Fluke Graphical Multimeter, 
	CE

	Microcal  
	CE

	Oxygen BOC Medical 99.5% License 07351
	

	Helium  BOC Medical 99.5% License 07351
	


	Test with Dry Air & 100% oxygen    Sensor S/n    
	615665

	100% Oxygen
	100%
	
	

	Dry Air 20.9
	20.8%
	
	

	Test with  Microcal
	DVM Reading
	Oxygen Target
	Oxygen Reading +/-1%

	Set Microcal to  +/- .05mV
	
	
	

	40mV
	39.9mV
	100
	100.0%

	30mV
	30.0mV
	75
	75.0%

	20mV
	20.0mV
	50
	50.0%

	10mV
	9.9mV
	25
	24.9%

	5mV
	4.9mV
	12.5
	12.3%

	1mV
	0.9mV
	2.5
	2.2%


Vandagraph Calibration Cert O2 Monitors 000000                            25 June 2014

