VANDAGRAPH

O2 Analyser
VN 202 - Repairs

VM3/COP/40.01
	Date: 23 Sep. 03
	Revision Date:15-Mar-06
	Issue: 1



	Test and Equipment Required


1.0
Visual and functional test required – EQUIPMENT REQUIRED?

	Goods-In


1.0 Book the goods into the system; the procedure is the same as Viamed Except that the complete packaging is forwarded to the Vandagraph office.

2.0
File a copy of the work sheet with the original customer correspondence on Repairs file and send a copy of both back with the repaired item.

	Method


1.0 Check the case and the cable for damage.

2.0 If the battery is in – check the battery condition, if not up to standard replace with a new battery.

3.0 Check the condition of the battery connections.

4.0 Check the L.C.D. for obvious faults.

5.0 Check the “Offset” at  “-3 to +3 (Maximum), if the reading is greater than “(1” then reset to optimum.

6.0 Add Sensor Check calibration in Air – this should be reading “20.9” steady. Check in 100% Oxygen, check calibration returns to “21%” in Air.
7.0 Check the cable Analyser end for intermittent faults, Check the cable at the Sensor end for intermittent faults.
8.0 Rotate the sensor and check for dead point.
9.0 Check auto – switch off.
i)
Yellow button switch-off automatic after 2-3 minutes



ii)
Green button push for “ON” push for “OFF”

10.0 Check the labels and replace where necessary.
11.0 Clean the case with a soft damp cloth, clean or replace the carrying case if necessary.

12.0 Fill in the work sheet completely.









VM3COP_40.01 VN202 Repairs           Page 1 of 1

