	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Master Blank

	R-33 DE Assembly

VM3/COP/40.13

	Date: 10-Mar-2006
	Revision Date: 17-May-07
	Issue 1

	Only to be attempted by trained personel





	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R33D1 Sensor
	
	

	1
	White Box & Lid
	
	

	1
	JST SM02B-SRSS-TB socket  (515-1490)
	
	

	2
	wire 1 cm from reel
	
	

	
	Super glue
	
	

	1
	Packing Material
	
	

	1
	Round “Vandagraph” Label
	
	

	2
	Serial number Labels
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	

	
	
	
	


	Assembly: Set Temperature of Iron to 350 C Fine bit


[image: image1.jpg]


[image: image2.jpg]



1. [image: image3.jpg]


Remove the R33D! Sensor from the Gas barrier bag, and carefully remove by unsoldering the Red & Black . .
Do not apply too much heat
2. Hold the sensor in a Flow divertor which is held in a vice.

3. [image: image4.jpg]


Scrape off the lacquer between the two solder points so that they will lie halfway down the connector

4. Glue the JST socket onto the PCB with Super glue Hold down until dry. 

5. [image: image5.jpg]


Ensure it is the correct way up
The rear terminals should point down towards the PCB And the narrowest part of the connector should be glued to the PCB








6. Strip the insulation off the wire for about 2.5cms.

7. Tin the wire for the full length

8. Solder the wires to the outer holes where the wires were removed and then to the socket 

9. Test the Sensor .The output should be between 23mV and 27mV. An Analox analyser can be used.




10. Remove the existing sensor label. 

11. Write on the Labels  R-33DE   and horizontally the Serial number & Date code.  
Or use printed labels with Citizen printer

12. Carefully remove the paper backing and stick the transparent cover down 

13. Carefully mount one label on the sensor and one label on the white container so that the edges of the labels are parallel to the edges of the container and Sensor.



14. Place packing material in the bottom and top of the white container and add the cap. Ensure it is sealed.

15. [image: image6.jpg]


[image: image7.jpg]


[image: image8.jpg]


[image: image9.jpg]


Add the Vandagraph round label to the top of the cap. Add the instruction leaflet, and place the whole into the polythene bag, seal it and affix a Serial Number Label.










VM3COP40.13 R-33DE Sensor Assembly 060924.doc        5/17/2007         Page 1 of 2

