	VANDAGRAPH
  COMPANY OPERATING PROCEDURE

	Oxygen Sensor manufacture

	R-22OMS sensor

VM3/COP/40.16

	Date: 30-Sep-03
	Revision Date: 17-May-07
	Issue 1

	Only to be attempted by trained personel





	Parts Required

	Quantity
	Description
	Part No.
	

	1
	R-22HDO
	
	

	1
	White Box 
	9030000
	

	1
	White Box Lid
	9030001
	

	1
	3way molex  rt angle
	
	

	1
	Packing Material
	
	

	1
	Round “Vandagraph” Label
	
	

	2
	Serial number Labels
	
	

	1
	Serial number Label for Outer bag
	
	

	1
	Polythene Bag (outer)
	
	


Warning  Use Soldering Iron set at 300C Do Not overheat the PCB on the Sensor
Procedure

1. Remove the R22HDO from the Gas barrier bag, 
2. Test the sensor 25mV +/- 30% 
3. Carefully remove the Molex plastic shroud

4. Unsolder the three pins

5. Solder into the holes the 3 way right angle Molex

[image: image1.jpg]



[image: image2.jpg]


[image: image3.jpg]


[image: image4.jpg]


[image: image5.jpg]



VM3COP40.16 R22OMS Sensor manufacture.doc       17/05/2007      Page 1 of 1

