VANDAGRAPH

          O2 Sensors

Teledyne R17D & R17VAN

Oxygen Sensor Test Procedure

VM3/COP/40.35
	Date: 5-Nov-03
	Revision Date: 17-Jul-06
	Issue: 1



	Test & Equipment Required


Functional test required using R17 tester and R17 test adapter,

(digital multimeter and 3.5mm jack plug adapter).

	Method


1. Open the packaging and remove the O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Connect the O2 sensor to the R17 tester, select 200mV range (marked with arrows) and wait for meter reading to stabilize.

4. Ensure that meter reading lies within the output specification range, 7mV-13mV
5. Disconnect the O2 sensor from the R17 tester and re-label if appropriate.

6. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated sticker and re-box.

7. Ensure the O2 sensor is booked out in the stock books.
(*).  
If the O2 sensor fails the output specification test, it should be resealed in it’s original packet and re-boxed.  It should then be placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in.









VM3COP40.35 R17 Oxygen Sensor 050204                               7/17/2006              Page 1 of 1

