VANDAGRAPH

          O2 Sensors

Teledyne R22D
Oxygen Sensor Test Procedure

VM3/COP/40.37
	Date: 5-Nov-03
	Revision Date: 17-Jul-06
	Issue: 1



	Test & Equipment Required


Functional test required using R22/R24 tester and R22/R24 test lead,

(digital multimeter and 3 pin molex adapter).

	Method


1. Open the packaging and remove fuel cell.

2. Check the O2 sensor for damage and signs of leaking electrolyte.

3. Connect the O2 sensor to the R22 tester, select 200mV range (marked with arrows) and wait for meter reading to stabilize.

4. Ensure that meter reading is within the output specification range, 7mV-13mV
5. Disconnect the O2 sensor from the R22 tester and re-label if appropriate.

6. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated sticker and re-box.

7. Ensure the O2 sensor is booked out in the stock books.
(*).  
If the O2 sensor fails the output specification test, it should be resealed in it’s original packet and re-boxed.  It should then be placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in.









VM3COP40.37 R22D Oxygen Sensor 060717     7/17/2006                             Page 1 of 1

