## Tabelle1
Release Date: | 2023-10-04 | [11 empty] | 2) | Detailed actions that should follow are based on the Risk Analysis of the HOST system.
Rev. | 4 | [11 empty] | 3) | Typical applications are:
Languages: | EN, DE (Parts) | [12 empty] | AM - Alarm monitor
Change: | update to protocol rev.9 (add TN-3.35.1, TN-3.35.2) | [12 empty] | SpC - Spot check measurement
Protocol: | [R1] SMARTsat_Communication_Protocol_rev.16 | [12 empty] | PTT - Sleep diagnostic/ PTT measurement
Intigration Guide: | [R2] SMARTsat_I-III_Integration_Guide_rev.11 | [12 empty] | All - All applications listed here
Protocol Viewer: | [R3] SMARTsatProtocViewer_v53.0.0.0e
Emulator: | [R4] SMARTsatEmulator_v14.0.0.1
FW: SMARTsat OEM III | B40.A27
FW: SMARTsat OEM II/I | A18.A25
[15 empty] | yes / no |  | pass/ fail/ NA
Test. Nr. | Channel | Ch. Name | ID | ID. Name | Byte | Bit | Item | Description and suggested Host action | Beschreibung und Host Aktion Vorschlag (German Translation, only mandatory parts) | Suggested Host Message | min. priority at Host Alarm | Typical Appl. 3) | Manda- tory | Test 1), 2) | IMPL | Description of implementation | Result
TN-1.1 | 0x01 | Common device | 0x01 | Protocol version | n/a | n/a | n/a | Information for service staff. Display version e.g. at the start-up screen or in the service menu. | N/A | N/A | N/A | All | NO | Request the protocol version at the EUT and verify that the received string equals the one set at the emulator.
TN-1.2 | 0x01 | Common device | 0x02 | Device identification | n/a | n/a | n/a | Information for service staff. Display version e.g. at the start-up screen or in the service menu. | N/A | N/A | N/A | All | NO | Request the device identification at the EUT and verify that the received string equals the one set at the emulator.
TN-1.3 | 0x01 | Common device | 0x03 | Firmware version | n/a | n/a | n/a | Information for service staff. Display version e.g. at the start-up screen or in the service menu. | N/A | N/A | N/A | All | NO | Request the firmware version at the EUT and verify that the received string equals the one set at the emulator.
TN-1.4 | 0x01 | Common device | 0x04 | Hardware version | n/a | n/a | n/a | Information for service staff. Display version e.g. at the start-up screen or in the service menu. | N/A | N/A | N/A | All | NO | Request the hardware version at the EUT and verify that the received string equals the one set at the emulator.
TN-1.5 | 0x01 | Common device | 0x05 | Serial number | n/a | n/a | n/a | Information for service staff. Display version e.g. at the start-up screen or in the service menu. | N/A | N/A | N/A | All | NO | Request the serial number at the EUT and verify that the received string equals the one set at the emulator.
TN-1.6 | 0x01 | Common device | 0x06 | Identifier sent once at device start-up | n/a | n/a | n/a | Repeated sending of this identifier indicates a repeated reset of the module. In this case display a technical message. A typical reason for reset of the module is an unstable or fluctuating supply voltage. | N/A | SpO2 Module defective / Device defective | low | All | YES | Set module switched on at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-1.7 | 0x01 | Common device | 0x06 | Identifier sent once at device start-up | n/a | n/a | n/a | Repeated sending of this identifier indicates a repeated reset of the module. In this case display a technical message. A typical reason for reset of the module is an unstable or fluctuating supply voltage | Wiederholtes Senden des IDENTEFIERS zeigt wiederholtes Neustarten des SMARTsat Moduls an. Zeige eine technische Fehlermeldung am Host System. Ein typischer Grund für den wiederholten Neustart von SMARTsat ist eine fluktuierende Spannungsversorgung | SpO2 Module defective / Device defective | low | All | YES | Set module switched on at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.1 | 0x02 | Error | 0x01 | Unknown channel | n/a | n/a | n/a | HOST should retry to send his request. Display a technical error message for at least 1sec upon occurrence, if the error continues. In addition check if the changed setting has been implemented (Ch=0x10: ID=0x04 Byte[6]). Also display the message, if the error occurs, independent of requests sent by the HOST. | Wiederholen sie die Anfrage an SMARTsat. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System, wenn sich die Fehlermeldung wiederholt. Prüfen sie ob ihre Anfrage übernommen wurde (Ch=0x10: ID=0x04 Byte[6]). Zeige auch eine Fehlermeldung an, wenn dieser Fehler gesendet wird unabhängig von ihrer Anfrage. | SpO2 Module defective / Device defective | low | All | YES | Set unknown channel at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.2 | 0x02 | Error | 0x02 | Unknown identifier | n/a | n/a | n/a | HOST should retry to send his request. Display a technical error message for at least 1sec upon occurrence, if the error continues. In addition check if the changed setting has been implemented (Ch=0x10: ID=0x04 Byte[6]). Also display the message, if the error occurs, independent of requests sent by the HOST. | Wiederholen sie die Anfrage an SMARTsat. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System, wenn sich die Fehlermeldung wiederholt. Prüfen sie ob ihre Anfrage übernommen wurde (Ch=0x10: ID=0x04 Byte[6]). Zeige auch eine Fehlermeldung an, wenn dieser Fehler gesendet wird unabhängig von ihrer Anfrage. | SpO2 Module defective / Device defective | low | All | YES | Set unknown identifier at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.3 | 0x02 | Error | 0x03 | Invalid parameter | n/a | n/a | n/a | HOST should retry to send his request. Display a technical error message for at least 1sec upon occurrence, if the error continues. In addition check if the changed setting has been implemented (Ch=0x10: ID=0x04 Byte[6]). Also display the message, if the error occurs, independent of requests sent by the HOST. | Wiederholen sie die Anfrage an SMARTsat. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System, wenn sich die Fehlermeldung wiederholt. Prüfen sie ob ihre Anfrage übernommen wurde (Ch=0x10: ID=0x04 Byte[6]). Zeige auch eine Fehlermeldung an, wenn dieser Fehler gesendet wird unabhängig von ihrer Anfrage. | SpO2 Module defective / Device defective | low | All | YES | Set invalid parameter at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.4 | 0x02 | Error | 0x04 | Selected baud rate is too slow | n/a | n/a | n/a | Raw Pleth (RP) / (RP2) need a minimum baud rate of 115200 Bd. Increase baud rate if RP /RP2 is selected. Display a technical error message for at least 1 sec upon occurrence, if the error continues. | Raw Pleth (RP) / (RP2) benötigt eine Mindestbaudrate von 115200 Bd. Erhöhen Sie die Baudrate, wenn RP /RP2 ausgewählt ist. Bei Auftreten soll das Host system mindestens s lang eine technische Fehlermeldung anzeigen, wenn der Fehler weiterhin besteht. | SpO2 Module defective / Device defective | low | All | YES | Set selected baud rate is too slow at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.5 | 0x02 | Error | 0x05 | Receive buffer overflow | n/a | n/a | n/a | Too much input data to module. Reset module upon repeated sending of the error. Display a technical error message for at least 1 sec upon occurrence, if the error continues. | Zu viele Daten werden an SMARTsat gesendet. Führe einen SW Reset durch, wenn SMARTsat diese Fehlermeldung sendet. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System, wenn sich die Fehlermeldung wiederholt. | SpO2 Module defective / Device defective | low | All | YES | Set receive buffer overflow at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.6 | 0x02 | Error | 0x06 | Frame corrupt, CRC error | n/a | n/a | n/a | HOST should retry to send his request. Display a technical error message for at least 1 sec upon occurrence, if the error continues. In addition check if the changed setting has been implemented (Ch=0x10: ID=0x04 Byte[6]). Also display the message, if the error occurs, independent of requests sent by the HOST. | Widerholen sie die Anfrage an SMARTsat. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System, wenn sich die Fehlermeldung wiederholt. Prüfen sie ob ihre Anfrage übernommen wurde (Ch=0x10: ID=0x04 Byte[6]). Zeige auch eine Fehlermeldung an, wenn dieser Fehler gesendet wird unabhängig von ihrer Anfrage. | SpO2 Module defective / Device defective | low | All | YES | Set frame corrupt, CRC error at the emulator error channel and send to the EUT – verify that an appropriate HOST action follows.
TN-2.7 | 0x02 | Error | 0x07 | Sensor Error: Red LED defective | n/a | n/a | n/a | The red LED at the sensor is defective. NOTE: status information ‘Sensor defective’ (Byte[0] Bit: 1 in channel 0x10, identifier 0x01) is sent at the same time . | n/a | Replace sensor/ Sensor fault | low | All | NO | Set Red LED defective at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.8 | 0x02 | Error | 0x08 | Sensor Error: Infrared LED defective | n/a | n/a | n/a | The infrared LED at the sensor is defective. NOTE: status information ‘Sensor defective’ (Byte[0] Bit: 1 in channel 0x10, identifier 0x01) is sent at the same time . | n/a | Replace sensor/ Sensor fault | low | All | NO | Set Infrared LED defective at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.9 | 0x02 | Error | 0x09 | Sensor Error: Photodiode defective | n/a | n/a | n/a | The detector at the sensor is defective. NOTE: status information ‘Sensor defective’ (Byte[0] Bit: 1 in channel 0x10, identifier 0x01) is sent at the same time . Display error message for at least 1sec upon occurrence of the error | n/a | Replace sensor/ Sensor fault | low | All | NO | Set Photodiode defective at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.10 | 0x02 | Error | 0x10 | Boot error | n/a | n/a | n/a | The boot-test during module switch-on failed. Display a technical error message for at least 1sec upon occurrence, thereafter reset the module. | Ein Fehler wurde bei der Prüfung des Speichers beim Einschalten von SMARTsat erkannt. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System und führe anschließend einen SW Reset durch. | SpO2 Module defective / Device defective | low | All | NO | Set Short circuit at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.11 | 0x02 | Error | 0x11 | Self-test error | n/a | n/a | n/a | The self-test during module switch-on failed. Display a technical error message for at least 1sec upon occurrence, thereafter reset the module. | Ein Fehler wurde während des Selbsttests beim Einschalten von SMARTsat erkannt. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System und führe anschließend einen SW Reset durch. | SpO2 Module defective / Device defective | low | All | YES | Set Boot error at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.12 | 0x02 | Error | 0x12 | Buffer overflow | n/a | n/a | n/a | An internal buffer overflow error occurred. Display a technical error message for at least 1sec, thereafter reset the module. | Ein Fehler in einem interner Speicher von SMARTsat wurde erkannt. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System und führe anschließend einen SW Reset durch. | SpO2 Module defective / Device defective | low | All | YES | Set Self-test error at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.13 | 0x02 | Error | 0x13 | Command to switch on auto scaled plethysmogram is not accepted | n/a | n/a | n/a | Auto scaled plethysmogram (ASP) and RP/RP2 cannot be switched on simultaneously. In RP/RP2 mode the ASP is not available to ensure minimal time jitter between sending of samples. The error code will show up, if the host enables ASP in RP/RP2 mode. Switch of RP/RP2 before switching on the ASP. | Um sicherzustellen, dass ein konstanter und definierter Time-Jitter zwischen den RP/RP2 Paketen gewährleistet wird, können ASP und RP/RP2 nicht zeitgleich eingeschaltet werden. Schalten sie zunächst RP/RP2 aus, bevor sie ASP einschalten. | SpO2 Module defective / Device defective | low | All | YES | Set Buffer overflow at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.14 | 0x02 | Error | n/a | Error log | n/a | n/a | n/a | It is recommended that the host stores all errors in an internal service error log to simplify root cause analysis in case of defects | n/a | SpO2 Module communication error | N/A | PTT | YES | Set Command is not accepted at the emulator and send to the EUT – verify that an appropriate HOST action follows.
TN-2.15 | 0x02 | Error | 0x0A | Sensor Error: Short circuit | n/a | n/a | n/a | Short circuit occurred in the sensor cable or connector. Measurement is interrupted. Disconnect sensor to reset the error. NOTE: status information ‘Sensor defective’ (Byte[0] Bit: 1 in channel 0x10, identifier 0x01) is sent at the same time . | n/a | store in the service menu of the host | N/A | All | NO | Read out the Error Log stored within the Host system and verify that all errors generated before were stored in the log.
TN-3.1 | 0x10 | Data | 0x01 | Status information | Byte[0] | Bit 0 | Sensor disconnected | Inform user that no sensor is connected. | Der HOST soll den Anwender darüber informieren, dass kein Sensor angeschlossen ist. | Connect sensor / Sensor disconnected | low | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows.
TN-3.2 | 0x10 | Data | 0x01 | Status information | Byte[0] | Bit 1 | Sensor defective | Inform user that the connected sensor is defective | Der HOST soll den Anwender darüber informieren, dass ein defekter Sensor angeschlossen ist. | Replace defective sensor / Sensor defective | low | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows.
TN-3.3 | 0x10 | Data | 0x01 | Status information | Byte[0] | Bit 2 | Wrong sensor | Inform user that the connected sensor is not compatible with the SpO2 module | Der HOST soll den Anwender darüber informieren, dass ein nicht-kompatibler Sensor angeschlossen ist. | Replace incompatible sensor/ Sensor incompatible | low | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows.
TN-3.4 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 0 | Finger out | Inform user that the sensor is no longer at the measurement site | Der HOST soll den Anwender darüber informieren, dass der Sensor sich nicht mehr am Messort befindet (von dem Finger abgerutscht ist). | Sensor off patient | low | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows.
TN-3.5 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 1 | Searching for pulse | Inform user that if the condition continues, the user should check the sensor application site and patient vital signs | n/a | Pulse search, no values (- -) displayed | N/A | All | NO | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.6 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 2 | Pulse searching longer than 30sec | Inform user that pulse search continues for a duration longer than 30 sec. User should check sensor application site and patient vital signs | Der HOST soll den Anwender darüber informieren, dass für eine Dauer von >30s kein Puls gefunden wurde. Es ist erforderlich den Messort sowie die Vitalfunktionen des Patienten zu überprüfen. | Pulse search longer than 30sec | message mandatory (80601-2-61 clause 201.12.4.101) | AM, CM | YES | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.7 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 3 | Low pulsation strength (IR AC/DC ratio <1%) | Inform user that the perfusion index (PI) is low. User should check the sensor application site and patient vital signs. Warming the application site improves perfusion. | n/a | Low perfusion | N/A | All | NO | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.8 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 4 | Low transmission | Inform user that the transmission signal at the sensor is low. Suggest to reposition the sensor, e.g. to a thinner finger | n/a | Low transmission, reposition sensor | N/A | All | NO | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.9 | 0x10 | Data | 0x01 | Status information | Byte[1] | Bit 7 | Loss of pulse | Suggests the user to check the sensor application site and patient vital signs because a sudden drop in perfusion was detected followed by pulse search. The flag is also set if for a duration longer than 5sec no measurement value was transmitted. This condition occurs during prolonged bad signal quality at the sensor site. | Der HOST soll den Anwender darüber informieren, dass es erforderlich ist den Messort sowie die Vitalfunktionen des Patienten zu überprüfen. Ein abrupter Abfall der Perfusion gefolgt durch eine erfolglose Pulssuche wurde detektiert. Dieser Status wird auch gesetzt, wenn für eine Dauer länger als 5sec kein Messwert übermittelt wird. Dieser Zustand tritt üblicherweise bei schlechter Signalqualität am Messort auf. | Loss of pulse/ Bad signal quality, (No value is displayed (- -)) | Alarm monitors should give at least medium priority alarm | AM, CM | YES | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.10 | 0x10 | Data | 0x01 | Status information | Byte[2] | Bit 0 | Ambient light | Inform the user that ambient light is distorting the measurement. Suggest covering sensor site. | Der HOST soll den Anwender darüber informieren, dass Umgebungslicht die Messung stört und dadurch die Messfunktion unterbricht. Das Abdecken des Messortes kann die Messfunktion erhalten. | Too much ambient light | low | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.11 | 0x10 | Data | 0x01 | Status information | Byte[2] | Bit 1 | Interferences detected | Inform the user that the sensor signal is disturbed by interferences (e.g. EMI). These affects the measurement - locate and isolate interference source | n/a | Signal interferences | N/A | All | NO | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.12 | 0x10 | Data | 0x01 | Status information | Byte[2] | Bit 2 | Motion artifacts | Inform the user that the sensor signal is disturbed by motion artifacts. Reduce motion at the sensor application site to increase measurement performance. | n/a | Motion artifacts | N/A | All | NO | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.13 | 0x10 | Data | 0x01 | Status information | Byte[2] | Bit 3 | Vital parameter out of range | Inform the user that the measurement is invalid due to abnormal signal properties e.g. intravascular dyes etc. | Der HOST soll den Anwender darüber informieren, dass die Messwerte außerhalb des Messbereiches liegen (z.B. bei Anwendung von Intravaskulären Farbstoffen) | Parameter out of range | N/A | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows. Verify IFU.
TN-3.14 | 0x10 | Data | 0x01 | Status information | Byte[2] | Bit 4 | Supply voltage out of range | Inform the user that the measurement accuracy is affected because the module supply voltage is out of range. Display a technical alarm message for at least 1 second upon occurrence. | Der HOST soll den Anwender darüber informieren, dass die Messgenauigkeit nicht gewährleistet ist, weil die Versorgungsspannung von dem Modul außerhalb des Messbereiches liegt. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System. | Device defective, send device to service department | medium | All | YES | Set status flags at the emulator and verify that an appropriate HOST action follows.
TN-4.01 | 0x10 | Data | 0x02 a | Auto scaled Plethysmogram (ASP) | n/a | n/a | n/a | The auto scaled plethysmogram is derived from the Infrared LED transmission signal. Display the plethysmogram at the host to visualize the signal and heart pulsation activity to the clinical staff. The plethysmogram may be switched off to extend battery life in case of low power host applications which clinically do not require display of a plethysmogram (see Ch=0x10: ID=0x18). | n/a | n/a | n/a | All | NO | Set Pleth ON at the emulator. Verify the correct display of the simulated emulator plethysmogram waveform at the EUT.
TN-4.02 | 0x10 | Data | 0x02 b | Pulse beep indicator | n/a | n/a | n/a | The pulse beep indicator can be used to emit a sound at the host correlating with the pulse rate frequency. It tags each detected pulse. In case no pulse tone sounds, the signal quality is not sufficient to reliably detect the single pulse as typically found under signal interference conditions (EMI, motion conditions) | n/a | n/a | n/a | All | NO | Monitor with sound functionality: Activation the Pulse Tone checkbox at the emulator. Verify that the EUT emits a periodic audible tone at the frequency of the pulse rate. Deactivate the Pulse Tone checkbox. Verify that the EUT does not emit a periodic audible tone.
TN-5.01 | 0x10 | Data | 0x03 | Raw Plethysmogram (RP) | n/a | n/a | n/a | Select RP or RP2 depending on monitor function. | n/a | n/a | n/a | PTT | NO | Set RP/ RP2 and sample Rate (SR) at the emulator. Verify the correct display of the simulated RP/ RP2 plethysmogram waveform at the EUT and implemented calculations (e.g. PTT).
TN-6.01 | 0x10 | Data | 0x04 | Results and indicators | Byte[0] | n/a | SpO2 Value [1…100 %], 0xFF = no value | Continuously display the SpO2 value, if no value is transmitted display “- -“. In case of single spot-check measurement, display value after 5 -10 seconds of measurement (also refer to response time modes identifier 0x10). | Der HOST soll die Messwerte kontinuierlich anzeigen. Wird kein Wert ausgegeben soll "--" angezeigt werden. Wird eine Spot-Check Messung durchgeführt, soll der Wert nach 5 bis 10 Sekunden erst angezeigt werden, wenn der Patient zur Ruhe gekommen ist (beachten sie zudem die möglichen Einstellungen der Messwert-Reaktionszeit). | n/a | n/a | All | YES | Cycle through valid SpO2 values (e.g. 100 to 0 %) using the decrease SpO2 functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-6.02 | 0x10 | Data | 0x04 | Results and indicators | Byte[1-2] | n/a | Pulse rate [x…y bpm], 0xFFFF = no value, (Hi+Lo) x = 20 bpm, EPR on x= 30 bpm, EPR off y = 300 bpm, EPR on y= 240 bpm, EPR off | Continuously display the pulse rate value, if no value is transmitted display “- -“. In case of single spot-check measurement, display value after 5 -10 seconds of measurement (also refer to response time modes identifier 0x10). | Der HOST soll die Messwerte kontinuierlich anzeigen. Wird kein Wert ausgegeben soll "--" angezeigt werden. Wird eine Spot-Check Messung durchgeführt, soll der Wert nach 5 bis 10 Sekunden erst angezeigt werden, wenn der Patient zur Ruhe gekommen ist (beachten sie zudem die möglichen Einstellungen der Messwert-Reaktionszeit). | n/a | n/a | All | YES | Cycle through valid pulse rate values (e.g. 300 to 20 BPM) using the decrease BPM functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC) – additionally verify that the frequency of the plethysmogram is changed accordingly during the test
TN-6.03 | 0x10 | Data | 0x04 | Results and indicators | Byte[3-4] | n/a | Perfusion Index PI= IAC/IDC [1...200 ‰] 0xFFFF = no value, a) Format with identifier 0x1D > 0x01 = resolution of PI is 0.1% b) Format with identifier 0x1D > 0x02 = resolution of PI is 0.01% | If a bar graph is implemented the perfusion index can be used to continuously define the bar graph amplitude during measurement. Due to the nature of the PI a logarithmic correlation is suggested . It is recommended to change the bar graph amplitude with changing signal quality. | n/a | n/a | n/a | All | NO | Cycle through valid pulsation strength values (e.g. 20.0 to 0.1 %) using the decrease pulsation strength functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-6.04 | 0x10 | Data | 0x04 | Results and indicators | Byte[5] | n/a | Signal quality [1...100 %], 0xFF = no value | It is mandatory to implement an indicator to show signal inadequacy when applicable (IEC 80601-2-61 (201.12.4.102)). Use at least symbol ISO 7000-0435 if signal quality below 61 %. If possible it is recommended to implement a bar graph representing the plethysmogram movement (based on channel 0x10, identifier 0x02). The color of the bar graph can be used as indicator for signal quality and measurement performance. 100 - Green: Good signal quality resulting in accurate measurement values. 60 %- Yellow: Reduced signal quality, which may reduce measurement accuracy. Also applies if following flags are set: “Supply voltage out of range”, “Vital parameter out of range” 30 %- Red: Poor signal quality, measurement values may be inaccurate 1% - Red: No SpO2 and PR measurement values available | Gemäß IEC 80601-2-61 (201.12.4.102) ist es verpflichtend den Anwender bei Messwerten die gegebenenfalls unzuverlässig /außerhalb der spezifizierten Messgenauigkeit liegen, zu informieren. Der HOST muss daher mindestens Symbol ISO 7000-0435 anzeigen, wenn die Signalqualität unterhalb 61% sinkt. Es ist empfohlen einen Bargrafen zu implementieren, welcher sich in Abhängigkeit der Pulsamplitude auf und ab bewegt (basierend auf Ch=0x10, ID=0x02). Dabei soll sich die Farbe des Bargrafen abhängig von der Signalqualität ändern: 100 - 61% - Grün: sehr gute Signalqualität, hohe Messgenauigkeit 60 - 31% - Gelb: reduzierte Signalqualität, Messwerte können außerhalb der spezifizierten Messgenauigkeit liegen 30 - 1% - Rot: sehr schlechte Signalqualität, Messwerte sind nicht zuverlässig | n/a | n/a | All | YES | Cycle through valid signal quality values (e.g. 100 to 1 %) using the decrease signal quality functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-6.05 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 0 | Response time: stable | Check continuously the settings of SMARTsat® by the host especially in case of continuous monitoring applications in which the HOST changes at start-up the default settings at the SMARTsat® module. Reason: In case SMARTsat® resets (e.g. fluctuating supply voltage), the setting returns to default. The HOST must set the HOST specific settings again if the module was reset. See also: Common device channel 0x01, 0x06 to identify a module restart. | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.06 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 1 | Response time: standard | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.07 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 2 | Response time: sensitive | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.08 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 3 | Response time: 8-Beat averaging | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.09 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 4 | Response time: 4-Beat averaging | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.10 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 5 | Pulse Rate mode: Standard mode (30 – 240 bpm) | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.11 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 6 | Pulse Rate mode: EPR Mode (20 – 300 bpm) | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-6.12 | 0x10 | Data | 0x04 | Results and indicators | Byte[6] | Bit 7 | New measurement (SpO2 and pulse rate value new. The longest Data Update Period is 28 sec) | This feature is useful during continuous monitoring to indicate periods in which the measurement value is not updated due to signal distortion. An indicator may be set at the host to show the user that data are not updated (e.g. circle element is filled proportional to fraction of 28sec) NOTE: if the distortion period exceeds 28sec, measurement values are set to no value “- -“ | n/a | n/a | n/a | AM, CM, PTT | NO | Uncheck the Value calc. new checkbox. Verify host action and IFU.
TN-7.01 | 0x10 | Data | 0x05 | Results and indicators | Byte[0-1] | n/a | SpO2 Value [0.01…100 %], 0xFF = no value | Continuously display the SpO2 value, if no value is transmitted display “- -“. In case of single spot-check measurement, display value after 5 -10 seconds of measurement (also refer to response time modes identifier 0x10). | Der HOST soll die Messwerte kontinuierlich anzeigen. Wird kein Wert ausgegeben soll "--" angezeigt werden. Wird eine Spot-Check Messung durchgeführt, soll der Wert nach 5 bis 10 Sekunden erst angezeigt werden, wenn der Patient zur Ruhe gekommen ist (beachten sie zudem die möglichen Einstellungen der Messwert-Reaktionszeit). | n/a | n/a | All | YES | Cycle through valid SpO2 values (e.g. 100 to 0.01 %) using the decrease SpO2 functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-7.02 | 0x10 | Data | 0x05 | Results and indicators | Byte[2-3] | n/a | Pulse rate [x…y bpm], 0xFFFF = no value, (Hi+Lo) x = 20 bpm, EPR on x= 30 bpm, EPR off y = 300 bpm, EPR on y= 240 bpm, EPR off | Continuously display the pulse rate value, if no value is transmitted display “- -“. In case of single spot-check measurement, display value after 5 -10 seconds of measurement (also refer to response time modes identifier 0x10). | Der HOST soll die Messwerte kontinuierlich anzeigen. Wird kein Wert ausgegeben soll "--" angezeigt werden. Wird eine Spot-Check Messung durchgeführt, soll der Wert nach 5 bis 10 Sekunden erst angezeigt werden, wenn der Patient zur Ruhe gekommen ist (beachten sie zudem die möglichen Einstellungen der Messwert-Reaktionszeit). | n/a | n/a | All | YES | Cycle through valid pulse rate values (e.g. 300 to 20 BPM) using the decrease BPM functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC) – additionally verify that the frequency of the plethysmogram is changed accordingly during the test
TN-7.03 | 0x10 | Data | 0x05 | Results and indicators | Byte[4-5] | n/a | Perfusion Index PI= IAC/IDC [1...200 ‰] 0xFFFF = no value, a) Format with identifier 0x1D > 0x01 = resolution of PI is 0.1% b) Format with identifier 0x1D > 0x02 = resolution of PI is 0.01% | If a bar graph is implemented the perfusion index can be used to continuously define the bar graph amplitude during measurement. Due to the nature of the PI a logarithmic correlation is suggested . It is recommended to change the bar graph amplitude with changing signal quality. | n/a | n/a | n/a | All | NO | Cycle through valid pulsation strength values (e.g. 20.0 to 0.1 %) using the decrease pulsation strength functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-7.04 | 0x10 | Data | 0x05 | Results and indicators | Byte[6] | n/a | Signal quality [1...100 %], 0xFF = no value | It is mandatory to implement an indicator to show signal inadequacy when applicable (IEC 80601-2-61 (201.12.4.102)). Use at least symbol ISO 7000-0435 if signal quality below 61 %. If possible it is recommended to implement a bar graph representing the plethysmogram movement (based on channel 0x10, identifier 0x02). The color of the bar graph can be used as indicator for signal quality and measurement performance. 100 - Green: Good signal quality resulting in accurate measurement values. 60 %- Yellow: Reduced signal quality, which may reduce measurement accuracy. Also applies if following flags are set: “Supply voltage out of range”, “Vital parameter out of range” 30 %- Red: Poor signal quality, measurement values may be inaccurate 1% - Red: No SpO2 and PR measurement values available | Gemäß IEC 80601-2-61 (201.12.4.102) ist es verpflichtend den Anwender bei Messwerten die gegebenenfalls unzuverlässig /außerhalb der spezifizierten Messgenauigkeit liegen, zu informieren. Der HOST muss daher mindestens Symbol ISO 7000-0435 anzeigen, wenn die Signalqualität unterhalb 61% sinkt. Es ist empfohlen einen Bargrafen zu implementieren, welcher sich in Abhängigkeit der Pulsamplitude auf und ab bewegt (basierend auf Ch=0x10, ID=0x02). Dabei soll sich die Farbe des Bargrafen abhängig von der Signalqualität ändern: 100 - 61% - Grün: sehr gute Signalqualität, hohe Messgenauigkeit 60 - 31% - Gelb: reduzierte Signalqualität, Messwerte können außerhalb der spezifizierten Messgenauigkeit liegen 30 - 1% - Rot: sehr schlechte Signalqualität, Messwerte sind nicht zuverlässig | n/a | n/a | All | YES | Cycle through valid signal quality values (e.g. 100 to 1 %) using the decrease signal quality functionality in the emulator and verify that is takes less than two seconds to display all values (after value is changed at the PC)
TN-7.05 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 0 | Response time: stable | Check continuously the settings of SMARTsat® by the host especially in case of continuous monitoring applications in which the HOST changes at start-up the default settings at the SMARTsat® module. Reason: In case SMARTsat® resets (e.g. fluctuating supply voltage), the setting returns to default. The HOST must set the HOST specific settings again if the module was reset. See also: Common device channel 0x01, 0x06 to identify a module restart. | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.06 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 1 | Response time: standard | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.07 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 2 | Response time: sensitive | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.08 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 3 | Response time: 8-Beat averaging | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.09 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 4 | Response time: 4-Beat averaging | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.10 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 5 | Pulse Rate mode: Standard mode (30 – 240 bpm) | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.11 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 6 | Pulse Rate mode: EPR Mode (20 – 300 bpm) | see: Ch=0x10, ID=0x04, Byte[6], Bit 0 | n/a | n/a | n/a | AM, CM, PTT | NO | Change the mode at the emulator and verify the host action
TN-7.12 | 0x10 | Data | 0x05 | Results and indicators | Byte[7] | Bit 7 | New measurement (SpO2 and pulse rate value new. The longest Data Update Period is 28 sec) | This feature is useful during continuous monitoring to indicate periods in which the measurement value is not updated due to signal distortion. An indicator may be set at the host to show the user that data are not updated (e.g. circle element is filled proportional to fraction of 28sec) NOTE: if the distortion period exceeds 28sec, measurement values are set to no value “- -“ | n/a | n/a | n/a | AM, CM, PTT | NO | Uncheck the Value calc. new checkbox. Verify host action and IFU.
TN-8.01 | 0x10 | Data | 0x06 | Sensor type | n/a | n/a | n/a | SMARTsat® identifier which sensor type is connected. Byte[0-1]: sensor type value = Byte[0]*0x100 + Byte[1] The standard SMARTsat® sensors are listed below: 0xFFFF /65535= undefined sensor type (see status information in identifier 0x01for detailed reason e.g. sensor disconnected or wrong) 0x000A /10 = Type C Sensors (closed sensors): SC7500 SoftCap; SCM7500 SoftCap medium; SCP7500 SoftCap pediatric; SF7500 SoftFlap, BF7500 Baby Foot Spotcheck Sensor 0x0028 /40 = Type O Sensors (open sensors): W7500 SoftWrap; 10-AP Disposable adult; 10-PP Disposable pediatric; 10-IP Disposable infant, Y7500 Disposable Multisite Y Sensor 0x0033 /51 = Type E Sensors (ear sensors): EP7500 Ear Probe 0x005B /91 = Type N Sensors (neonatal sensors): 10-NP Disposable neonatal NOTE: Customer specific firmware with customized Sensor Type Codes are listed in a separate document | n/a | n/a | n/a | All | NO | Cycle through all available sensor types in the emulator and request the sensor type on the EUT – verify that the correct sensor type is identified by the EUT
TN-8.02 | 0x10 | Data | 0x10 | Response Time setting | n/a | n/a | n/a | The selected response time setting has influence on the motion resistance performance and the response time of SpO2 and PR measurement (see section 2.2.2 of [R2] for more detail). Depending on the intended use of the host a selection or all response time modes may be implemented. For sleep diagnostic 4-beat averaging is typically used, during continuous monitoring during daily activity the standard or stable mode is recommended. | n/a | n/a | n/a | All | NO | Request the response time setting on the EUT and verify that the data received equals the setting in the emulator.Cycle through and set all available response time settings on the EUT – verify that the received message in the emulator equals the setting on the EUT
TN-8.03 | 0x10 | Data | 0x11 | High Resolution Plethysmogram (HRP) | n/a | n/a | n/a | Deactivated since Host protocol rev.9 | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
TN-8.04 | 0x10 | Data | 0x12 | Pulse rate mode | n/a | n/a | n/a | The standard pulse rate mode covers the typical pulse rate measurement range in humans (30 - 240bpm). In case of veterinary applications it can be suitable to activate the enhanced pulse rate mode (20 - 300bpm). | n/a | n/a | n/a | All | NO | Request the pulse rate mode on the EUT and verify that the result received equals the setting in the emulator. Cycle through and set all available pulse rate modes on the EUT– verify that the received message in the emulator equals the setting on the EUT
TN-8.05 | 0x10 | Data | 0x17 | Set Send Frequency of status information | n/a | n/a | n/a | Typically the status information is sent at 5Hz. For low power handheld host applications the battery life is extended by reducing the data sent by SMARTsat®. This is mainly achieved by deactivating the RP/ RP2; in addition data rates can be reduced by reducing the send frequency of the status to 1 Hz. This has to be considered in the risk analysis depending on the intended use of the host application. | n/a | n/a | n/a | All | NO | Request the status information send frequency on the EUT and verify that the result received equals the setting in the emulator. Cycle through and set all available status information send frequencies on the EUT – verify that the received message in the emulator equals the setting on the EUT. Verify that the status information frames are received with the correct frequency by the EUT
TN-8.06 | 0x10 | Data | 0x18 | Set Auto scaled plethysmogram | n/a | n/a | n/a | see: Ch=0x10, ID=0x02 Refer to section 3.2 of the Integration Guide [R2] for more information. | n/a | n/a | n/a | All | NO | Request the auto scaled plethysmogram setting on the EUT and verify that the result received equals the setting in the emulator. Switch the auto scaled plethysmogram to OFF in the EUT and verify that no waveform is displayed. Switch the auto scaled plethysmogram to ON in the EUT and verify that the waveform is displayed.
TN-8.07 | 0x10 | Data | 0x19 | Switch HR pleth | n/a | n/a | n/a | Possibility to switch Raw Pleth on/off independent of sampling rate. Refer to section 2.2.3 of the Integration Guide [R2] and the host protocol [R1] for more information. NOTE: for host protocol rev.8 and smaller use the HRP function | n/a | n/a | n/a | All | NO | Functionality of the PTT feature is not part of this verification. Verify correct implementation of the protocol and verify that an error is send in case the PR/ PR2 is switched off at the emulator during use.
TN-8.08 | 0x10 | Data | 0x1A | Set sampling rate | n/a | n/a | n/a | Possibility to select the desired sampling rate. The measurement accuracy is increased with higher sampling rates, but also the power consumption is increased. The raw plethysmogram (RP/RP2) is transmitted in the selected sampling rate. | n/a | n/a | n/a | All | NO | Cycle through all settable sampling rates at the EUT and verify that EUT behaves as specified.
TN-8.09 | 0x10 | Data | 0x1B | Raw Plethysmogram (RP2) | n/a | n/a | n/a | Refer to section 2.2.3 of the Integration Guide [2] and the host protocol [1] for more information. | n/a | n/a | n/a | PTT | NO | Functionality of the PTT feature is not part of this verification. Verify correct implementation of the protocol and verify that an error is send in case the PR/ PR2 is switched off at the emulator during use.
TN-8.10 | 0x10 | Data | 0x1C | Set SpO2 results Resolution | n/a | n/a | n/a | 0x01 = send results with identifier 0x04 each second. (integer representation: SpO2 results without decimal digits) 0x02 = send results with identifier 0x05 each second (float representation: SpO2 results with 2 decimal digits) | n/a | n/a | n/a | All | NO | Set desired resolution and verify displayed data
TN-8.11 | 0x10 | Data | 0x1D | Set perfusion index (PI) Resolution | n/a | n/a | n/a | 0x01 = resolution of PI is 0.1% 0x02 = resolution of PI is 0.01% (default setting since protocol rev.12). Used for OEM I and II which perform at lower perfusion (down to 0.02%) | n/a | n/a | n/a | All | NO | Set desired resolution and verify displayed data
TN-8.12 | 0x10 | Data | 0x1F | Request current module settings | n/a | n/a | n/a | Used to verify that the correct setting is set at the SMARTsat module. Result may be used to show setting at the host system or to check if send command was accepted in case the SMARTsat® response was not received upon change of setting. | n/a | n/a | n/a | All | NO | Change externally adjustable parameter at the emulator. Request current setting in the host and verify that the host requests change of parameter to the correct initial settings before manipulation.
TN-8.13 | 0x10 | Data | 0x30 | Software reset | n/a | n/a | n/a | The host may reset the SMARTsat by sending a SW command, through the reset pin or by switching the power supply. Consider that upon restart the SMARTsat returns to the default settings. The settings selected by the user after power on should therefore be preserved after reset initiated by the host system without user action. Refer to section 4.3 of the Integration Guide [R2] for more information. | Siehe auch Error channel 0x02 | n/a | n/a | All | YES | Verify that the implemented HW reset (power on) and SW reset (command send by host and/ or reset pin) are functional. Ensure that selected user settings (externally adjustable parameter) are restored after reset by reset pin or software host command.
TN-8.14 | 0x10 | Data | 0x31 | Baud rate setting | n/a | n/a | n/a | For low power handheld host applications the battery life is extended by selecting the highest baud rate available at the SMARTsat®. Note that lower baud is also available for host systems with limited baud rate. Note that the selected baud rate is stored in the permanent flash. At each restart the last selected baud is valid. The baud can be changed maximal 1000 times. After this the baud rate cannot be changed again. | n/a | n/a | n/a | All | NO | Request the baud rate setting on the EUT and verify that the result received equals the setting in the emulator. Cycle through the available baud rates on the EUT and verify that valid data is received (e.g. auto scaled plethysmogram is displayed properly) after the baud rate change
TN-8.15 | n/a | n/a | n/a | communication errors | n/a | n/a | Frame counter error | Each frame has a continuously incremented number. By verifying increment of the counter, lost frames are detected. In case too many frames are lost the display and update of the measurement results is no longer guaranteed. Display a technical alarm message | Jedes Datenpacket wird durchnummeriert. Durch Verifizieren des Zählerinkrements werden verlorene Datenpakete erkannt. Bei zu großem Verlust von Daten, ist die Anzeige und Aktualisierung der Messergebnisse nicht mehr gewährleistet. Zeige eine technische Fehlermeldung mit Mindestdauer von 1s am Host System an. | Device defective / SpO2 Module communication error | n/a | All | YES | Provoke frame counter errors using the emulator – verify that an appropriate message or action is seen on the EUT
TN-8.16 | n/a | n/a | n/a | communication errors | n/a | n/a | Frame error | Communication and timing errors, including distortion by EMI, can result in partially received frames. These can potentially transport faulty data. Analyze the structure of each received data frame to ensure complete data transmission. Do not display or store data sent in a corrupted frame or frame with CRC error. Display a technical alarm condition upon frame errors | Kommunikations- und Timmingfehler, einschließlich Störungen durch EMI, können unvollständige Datenpakete zur Folge haben. Diese können möglicherweise fehlerhafte Daten transportieren. Analysiere die Struktur jedes empfangenen Datenrahmens, zur Gewährleistung einer vollständigen Datenübertragung. Beschädigte Datenpakete oder Pakete mit CRC-Fehlern dürfen nicht am HOST System angezeigt oder gespeichert werden. Zeige bei kontinuierlichem Auftreten der Fehler eine Technische Fehlermeldung mit Mindestdauer von 1s am Host System an. | Device defective / SpO2 Module communication error | n/a | All | YES | Provoke frame / CRC errors using the emulator – verify that an appropriate message or action is seen on the EUT
TN-8.17 | n/a | n/a | n/a | communication errors | n/a | n/a | Unknown data | Display a technical alarm condition to make the user aware that a communication error occurred. | Zeige bei kontinuierlichem Auftreten der Fehler eine technische Fehlermeldung mit Mindestdauer von 1s am Host System an, um den Anwender darauf aufmerksam zu machen, dass Kommunikationsfehler vorliegen. | Device defective / SpO2 Module communication error | n/a | All | YES | Provoke unknown data using the emulator – verify that an appropriate message or action is seen on the EUT

## History
Template HISTORY LOG
REV | Date | Changes
0 | 2016-12-05 | Initial version in Excel File format “APPENDIX 1_Host Validation Checklist_rev.0_EN_DE” Created together with documents and tools: SMARTsat_I-III_Integration_Guide_rev6; SMARTsat_Communication_Protocol_rev7-A; SMARTsatProtocViewer_v0.33e; SMARTsatEmulator_v12.7 FW OEM III: B19.A14
1 | 2017-02-20 | Add Identifier 0x0A (Sensor Error: Short circuit) to Error channel (TN-2.15) Created together with documents and tools: SMARTsat_I-III_Integration_Guide_rev6; SMARTsat_Communication_Protocol_rev8; SMARTsatProtocViewer_v0.34e; SMARTsatEmulator_v13.2 FW OEM III: BM.03.B27.A18.1X
2 | 2018-03-01 | Update to protocol rev.9: - Add Identifier 0x19 (Switch High Resolution Plethysmogram ON/OFF) TN-3.35.1 - Add Identifier 0x1A (Set sample rate), TN-3.35.2 Created together with documents and tools: SMARTsat_I-III_Integration_Guide_rev7; SMARTsat_Communication_Protocol_rev9; SMARTsatProtocViewer_v0.37e; SMARTsatEmulator_v13.2 FW OEM III: B33.A22; FW OEM I/ II: A15.A22
3 | 2020-05-20 | Update to protocol rev.11: - Add RP (Raw Pleth) Created together with documents and tools: SMARTsat_I-III_Integration_Guide_rev9; SMARTsat_Communication_Protocol_rev11; SMARTsatProtocViewer_v0.37e; SMARTsatEmulator_v13.2 FW OEM III: B33.A22; FW OEM I/ II: A15.A22
4 | 2023-10-04 | Add section Reference (2) Rename to “Protoco l Implementation Checklist” (before Host Validation Checklist) Update to Protocol rev.16: --identifier 0x05 (Float SpO2 results) - identifier 0x07 (Raw Plethysmogram output of two wave length) - identifier 0x1B (Switch Raw Plethysmogram output of two wave length (RP2) ON/OFF) - identifier 0x1C (Result output switch between identifier 0x04 and 0x05) - ASP with 1 data point at 75 Hz, (see Ch 10, Id 0x02, 2) - identifier 0x1D (resolution of perfusion index in results identifier 0x04) Created together with documents and tools: SMARTsat_I-III_Integration_Guide_rev11; SMARTsat_Communication_Protocol_rev16; SMARTsatProtocViewer_v53.0.0.0e; SMARTsatEmulator_v14.0.0.1 FW OEM III: B40.A27; FW OEM I/ II: A18.A25