[image: image52.emf]


CO Check Pro / CO Check Baby
Operating Manual

MD Diagnostics Ltd

15 Hollingworth Court
Turkey Mill

Ashford Road

Maidstone

Kent ME14 5PP
UK

Tel: + 44 (0) 1622 682686
Fax: + 44 (0) 1622 681693
Email: sales@mdd.org.uk
Website: http://www.mdd.org.uk
Revision 10.0 March 2016
CONTENTS

2Package Contents


4Overview


4Operation


7Warning/Cautions


7User Menu


7Calibrate


8Calibration


10Setting Environment Mode


10Setting Breath Hold Time


11Battery Life


12Power Saving


12Cleaning


13Servicing


13Spares


14Specifications


19Appendix A – CO Check and Smoking Cessation


19Appendix B – CO Check and Environmental Capture




Package Contents
The CO Check Pro / CO Check Baby devices are supplied with:

1. CO Check Pro (Cat No. CO20) / CO Check Baby (Cat No. CO30)
2. 22 mm reducing connector for calibration (Cat No. CSC01)

3. PP3 Battery (Cat No. PP3B)

4. 6 SafeBreath™ filtered cardboard mouthpieces (Box of 200 Cat No. FM200)
5. Hard shell plastic carry case (Cat No. CS101)

6. Operating Manual (Cat No. OPM01)

7. CO / Pregnancy CO - Guide charts (Cat No. GC01) / (Cat No. GC02)
[image: image2.png]—





 

[image: image3.png]


Keypad Interface  
[image: image4.png]


  = ON/OFF Power button
[image: image5.png]


  = Select Function button
* See cleaning guidance on page 11. 
Overview
The CO Check Pro / CO Check Baby are handheld portable battery-operated devices used for measuring the concentration of carbon monoxide (CO) in the breath, calculating the percentage of carboxyhaemoglobin (%COHb) in the blood and for the CO Check Baby, also the percentage of foetal carboxyhaemoglobin (%FCOHb).
They are designed to be used primarily as a screening tool for smoking cessation, but it can also be used by Accident & Emergency departments, GP surgeries and by the fire fighting service. In smoking cessation clinics, it is used for checking the subject’s progress and compliance. A&E clinicians, GPs and fire fighters can use the device to quickly assess the level of suspected CO poisoning.
They are easy to use, very accurate and require a single breath into the device to display CO results in parts per million (PPM) and the percentage of carboxyhaemoglobin, %COHb and the percentage of foetal carboxyhaemoglobin (%FCOHb) - (CO Check Baby only).
They also measure ambient CO on switch ON to ensure that the environment is free from CO contamination (if enabled).

Operation

Insert the 9V PP3 battery (#3 supplied) by removing the battery cover and clipping the battery in place, and then replace the battery cover. Insert the SafeBreath™ filtered cardboard mouthpiece (#4).

Turn the CO Check ON by pressing ON/OFF Power key  [image: image6.png]


 . The device will display the version number before starting the countdown timer. The unit will countdown for 10 seconds to ensure that the sensor is stabilised. When the countdown timer reaches zero it will display the current environment reading if enabled.

[image: image7.png]



If the Environmental monitor is disabled, it will zero to ambient air automatically. 
The subject should be encouraged to hold the breath for 15 seconds (CO Check Pro) or 10 secs (CO Check Baby). Press the Select Function key [image: image8.png]


  and the display will show the Blow icon:
[image: image9.png]



Alternatively, set the hold breath timer to desired number of seconds. The unit will count down the period during which the subject should hold the breath. When the period expires, the unit will display the blow icon. 
[image: image10.bmp]
[image: image11.png]


 
If the environment reading is disabled, the original countdown will be based on hold breath setting, provided it is more than 10 seconds.
The subject should place their lips around the cardboard mouthpiece and blow gently and continue blowing until their lungs are completely empty. CO is collected in the last portion of the breath (alveolar breath). 

Test reading display

The CO Check will beep and display the final reading in PPM (parts per million). By pressing the Select Function key[image: image12.png]


 , the reading scrolls to %COHb (carboxyhaemoglobin) and then finally to % FCOHb (foetal carboxyhaemoglobin) - (CO Check Baby only). Pressing the Select Function key [image: image13.png]


 again reverts back to PPM.
[image: image14.png]


 
The maximum reading the CO Check Pro can measure is 99 PPM. Any reading above that will be displayed as - - (over range).
Coloured lights are there to represent and provide a visual display of the resulting PPM reading. The coloured lights represent the following:
0 – 6 ppm 

= Green Light 


= Non Smoker

7 – 10 ppm 

= Yellow Light 


= Light Smoker

11 – 20 ppm 
= Red Light 



= Smoker

20+ ppm 


= Flashing Red Light 
= Heavy Smoker
For %FCOHb only:

Green light

= Healthy

Red light


= High Risk

[image: image15.png]


 
The Red Light indication for a smoker is designed to make an expectant mother realise the effect of her smoking on her unborn child. 

After examining the readings, the unit can either be switched off using the ON/OFF Power key [image: image16.png]


, or another measurement can be carried out by pressing the Select Function [image: image17.png]


 key for at least three seconds. 
[image: image18.png]


 
To save the battery, the unit will switch OFF automatically after three minutes if not used.
Warning/Cautions

· If higher than expected levels of CO are displayed this could be due to CO poisoning and medical attention should be immediately sort.

· The CO Check Pro has a cross-sensitivity to hydrogen (from some gastro-intestinal disorders) which may affect the reading.

· Cleaning with products that contain alcohol may result in permanent damage to the sensor. See Spares section on page 12 for non-alcohol wipes.
· The battery should be changed when the low battery icon appears on the display.
· Please only use accessories supplied by MD Diagnostics to ensure the device performs as intended.
The disposable cardboard mouthpieces are single-use only.  Re-use of these single use mouthpieces can increase the risk of cross infection and the mouthpiece should be disposed of after use by the patient.
User Menu

The menu allows the user to either calibrate the unit or set the environment mode option. To enter the User Menu, press ON/OFF Power key [image: image19.png]


  the unit with Select Function key [image: image20.png]


 pressed simultaneously. Do not release the Select Function key [image: image21.png]


 until the User Menu appears. 

User Menu






Calibrate





Environment Mode






Breath Hold Time






Set Light colours





Set Baby colours





Exit
The first option is highlighted. To move between options, press the Select Function key [image: image22.png]


 momentarily (< 0.5 S). To select the highlighted option, press the Select Function key [image: image23.png]


 for at least 3 seconds to confirm.
Calibration

The device should be recalibrated at least once every six months. Calibration gas can be purchased from MD Diagnostics Ltd, see Spares for full ordering details. 
[image: image24.png]Control

Valve 4

20 ppm
carbon monoxide
in air

Flow
Indicator

Plastic Tubing

T

b

Reducing Connector

for Calibration





To calibrate the device, perform the following steps:
1. Connect the control valve (Cat No. CV20) to the 20 ppm gas cylinder. Cat No. C20 calibration gas cylinder shown in the diagram above.
2. Connect the short plastic tubing reducing connector (Cat No CSC01).

3. Switch ON the device with the Select Function key [image: image25.png]


 pressed until the device displays ‘Calibration’. Release the Select Function key [image: image26.png]


.

4. The unit will countdown from 15 to 0 and then will display the calibration icon.

[image: image27.bmp]
5. Attach the connector and open the control valve (anti-clockwise) until the ball in the flow indicator is between the two marks. This will supply gas of approximately 0.25 L/min. 

6. Apply this flow until the device beeps after 20 seconds. The device will display the gas concentration reading.

7. If the reading is not 20 ppm, press the Select Function key [image: image28.png]


 for at least 3 seconds to accept the new calibration value. The device will say ‘done’ and then will show 20 ppm.

[image: image29.png]


 It is advisable to accept the calibration even though it displays 20 ppm (i.e. the ideal reading). This will reset the ‘calibration due’ timer.
[image: image30.png]


 Switch OFF the device if the calibration was incorrectly performed. Do not press the Select Function key [image: image31.png]


. 

· To prevent incorrect calibration, only the readings within the range of 16-24 can be calibrated. An ‘Error’ message will be displayed if the calibration is accepted for readings outside this range.

8. Switch OFF the device; wait for a minute and it is now ready for use.
If the device is not calibrated for six months, the device will beep three times before the countdown starts and will display the following message:
Calibration Expired
Please Calibrate
the unit
The reading is not
Guaranteed
The message will appear for three seconds and then the device will operate as normal. Calibration should be performed as soon as possible.
Setting Environment Mode

The Environment reading on start-up can be enabled or disabled. To change the setting, switch on the unit with Select Function key [image: image32.png]


  and enter the user menu. Select the ‘Environment Mode’ and it will display the options ‘On’ or ‘Off’. The current setting will be highlighted. Select the preferred option and press Select Function key [image: image33.png]


  for at least 3 seconds to confirm. The unit will restart with the new setting.

Setting Breath Hold Time

The breath hold time is only applicable in the Environment mode operation. After the environment reading, the unit will countdown the timer during which time the subject is encouraged to hold the breath. The hold time can be set from the following:



None



5 Seconds



10 Seconds



15 Seconds



20 Seconds



25 Seconds



30 Seconds

Use the Select Function key [image: image34.png]


   to scroll through the times, then press the Select Function key [image: image35.png]


  again for at least 3 seconds to confirm the required time. The unit will restart with the new setting.

Set Light colours

The following coloured warning lights which are a visual display of the resulting PPM reading, can be pre-set to illuminate at any value within the following ranges: 
Green 
(0 to 20ppm) 
Yellow 
(0 to 20ppm)
Red 

(0 to 30ppm)
Use the Select Function key [image: image36.png]


   to scroll through the values, then press the Select Function key [image: image37.png]


  again for at least 3 seconds to confirm the required value. The unit will restart with the new setting
Set Baby colours (applicable to CO Check Baby only)
The following coloured warning lights which are a visual display of the resulting foetal reading, can be pre-set to illuminate at any value within the following ranges: 
Green 
(0 to 20ppm)
Amber 
(0 to 20ppm)
Red

(0 to 30ppm)
Use the Select Function key [image: image38.png]


   to scroll through the values, then press the Select Function key [image: image39.png]


  again for at least 3 seconds to confirm the required value. The unit will restart with the new setting  

Battery Life

The 9V PP3 battery should provide at least 30 hours of continuous use. When the battery is low, the battery low icon will be displayed for three seconds when the device is switched on.
[image: image40.bmp]
The device can still be used, but it is advisable to replace the battery.

[image: image41.png]



It is recommended to use an Alkaline battery.

[image: image42.png]



The device will need recalibration after the replacement of the battery. Duracell notification - Alkaline batteries can be safely disposed of with normal household waste. Never dispose of batteries in fire because they could explode.
If the battery is very low such that the reading is not reliable, the device will display the ‘battery dead’ icon and will not operate until the battery is replaced.  

[image: image43.bmp]
Power Saving

To save battery power, the device will automatically turn itself OFF three minutes after the last key press. Do not remove the battery unless the device is not going to be used for a very long time. Remember, calibration will be required when the battery is reconnected.

[image: image44.png]


To prevent the device from switching OFF, press the Select Function key [image: image45.png]


 within three minutes of the last press or when the reading is first displayed.

Cleaning

The device can be cleaned using non-alcohol wipes (see Spares). Please be careful not to touch the surface of the sensor or allow moisture to do the same.
[image: image46.png]



Cleaning with products that contain alcohol may result in permanent contamination of the sensor.
Servicing
If your unit requires servicing or repair, then please contact your local authorised distributor or MD Diagnostics Ltd directly.

[image: image47.png]



The CO Sensor should be replaced every five years.
Email: 
service@mdd.org.uk

Tel: 

+ 44 (0) 1622 682686
Spares

	Catalogue Number
	Description


	C20
	20 litre can 20ppm CO balance air

	C110
	110 litre can 20ppm CO balance air

	CV20
	Control valve for 20 litre can

	CV110
	Control valve for 110 litre can

	FC50
	Fuel cell for CO Check Pro and CO Check Baby

	FM200
	SafeBreath™ 22mm Filtered Mouthpieces 200 per box

	MP500
	Standard 22mm Mouthpieces 500 per box

	MP250
	Eco Straw 10mm Mouthpieces 250 per box

	IP10
	Intermediate plastic adaptors 10 per pack for MP500

	IP20 
	Intermediate plastic adaptors 10 per pack for MP250

	TDW01
	Bacterial cleansing wipes 125 wipes per tub

	TDW12            
	Bacterial cleansing wipes 12 tubs per case

	PP3B
	9v PP3 alkaline battery

	CS100
	Replacement of fuel cell, calibration, battery and service

	GC01
	CO Guide chart (CO Check Pro only)

	GC02
	Pregnancy CO Guide chart (CO Check Pro Baby only)


Specifications
Gas Detected 






Carbon Monoxide

Concentration Range 



0-99 PPM

Detection Sensor Used 



Electrochemical fuel cell

Sensitivity 
0.1 PPM (0-10 PPM range); 1 PPM (10-99PPM range)
Accuracy (repeatability)



±2%
Operating Temperature



5-35 degrees Celsius

Operating Pressure 




Atmospheric 10%
Operating Humidity




30% to 90% RH
Operating Altitude




Sea level to 6000 ft

Storage Temperature




-20 to + 70O C

Storage Humidity





10% to 90% RH 
Hydrogen cross-sensitivity 

<12% at 20˚ C
Sensor Life 







2 - 5 years, 2 years’ warranty

Sensor Drift 






<2% per month

Display 








128 X 64-pixel Graphic LCD

Power Supply 






Single 9V PP3 battery

Weight (approximate)



160g including battery

Dimensions 






135mm x 65mm x 60mm

*Indicator levels 





Green 0 - 6 PPM












Yellow 7 - 10 PPM












Red 11 - 20 PPM












Flashing red 20+ PPM
* Also User configurable.
Symbols

[image: image1.jpg]SPNVD

DIAGNOSTICS LTD.




In accordance with Directive 93/42/EEC

0120

[image: image51.wmf]
Type B Device

[image: image48.jpg]



To avoid the potential effects on the environment and human health as a result of the presence of hazardous substances in electrical and electronic equipment, end users of electrical and electronic equipment should understand the meaning of the crossed-out wheeled bin symbol. Do not dispose of WEEE as unsorted municipal waste and have to collect such WEEE separately.
Important information regarding Electromagnetic Compatibility (EMC)

Medical devices may be susceptible to electromagnetic interference from other devices such as PC’s and mobile telephones. Electromagnetic interference may impair the operation of the medical device and could create a potentially unsafe situation.  

In order to regulate the requirements for EMC, to limit unsafe product situations, BS EN 60601-1-2 standard has been implemented. This standard defines the levels of Immunity to electromagnetic interference as well as the levels of electromagnetic Emissions for medical devices.  

As a medical device the CO Check conforms to BS EN60601-1-2 standard for both Immunity and Emissions.  
	Guidance and Manufacturer’s Declaration – Electromagnetic Immunity

	The CO Check is intended for use in the electromagnetic environment specified below.  The customer or the user should assure that it is used in such an environment.

	Immunity Test 


	IEC 60601 Test level
	Compliance level
	Electromagnetic environment guidance

	Radiated Immunity IEC 61000-4-3 


	10V/m
	10V/m
	Avoid use in environments likely to exceed 10V/m

	Electrostatic Discharge (ESD)  

IEC 61000-4-2 


	± 6 kV contact  

± 8 kV air 


	± 6 kV contact  

± 8 kV air 


	No restrictions in the intended environment

	Electrical fast transient/ burst  

IEC 61000-4-4 


	N/A
	N/A
	None

	Surge  

IEC 61000-4-5 


	N/A
	N/A
	None

	Voltage dips, short interruptions and voltage variations on power supply input lines  

IEC 61000-4-11
	N/A
	N/A
	None

	Power Frequency (50/60 Hz) magnetic field  

IEC 61000-4-8
	N/A
	N/A
	None

	NOTE:   UT is the a.c. mains voltage prior to application of the test level.


	Guidance and Manufacturer’s Declaration - Electromagnetic Emissions

	The CO Check is intended for use in the electromagnetic environment specified below.  The customer or the user should assure that it is used in such an environment. 

	Emissions Test
	Compliance level
	Electromagnetic environment guidance

	RF Emissions CISPR 11
	Group 1
	The CO Check uses RF energy only for its internal function.  Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment.

	RF Emissions CISPR 11
	Class B
	The CO Check can be used in domestic, light and heavy industrial environments. 

	Harmonic emissions IEC 61000-3-2
	[Not Applicable]
	

	Voltage fluctuations / flicker emissions IEC 61000-3-3
	[Not Applicable]
	

	
	[See 5.2.2.1 c) and Figure 1]
	The CO Check is suitable for use in all establishments, including domestic establishments and those directly connected to the public low-voltage power supply network that supplied buildings used for domestic purposes.

	
	[See 5.2.2.1 c) and Figure 1]
	The CO Check is suitable for use in all establishments other than domestic and those directly connected to the public low-voltage power supply network that supplied buildings used for domestic purposes.

	RF Emissions CISPR 14-1
	Complies
	The CO Check is not suitable for interconnection with other equipment. 

	RF Emissions CISPR 15
	Complies
	The CO Check is not suitable for interconnection with other equipment.


	Guidance and Manufacturer’s Declaration - Electromagnetic Immunity

	The CO Check is intended for use in the electromagnetic environment specified below.  The customer or the user should assure that it is used in such an environment. 

	Immunity Test
	IEC 60601 Test level
	Compliance level
	Electromagnetic environment guidance

	Conducted RF IEC 61000-4-6

Radiated RF IEC 61000-4-3
	3 Vrms 150 KHz to 80 MHz 

3 V/m 80 MHz to 2.5 GHz


	[V1] V  

[E1] V/m
	Portable and mobile RF communications equipment should be used no closer to any part of the CO Check, including cables, than the recommended separation distance calculated from the equation applicable to the frequency of the transmitter.  

Recommended separation distance

d = [ 3.5 ] √ P         

         V1
d = [3.5 ] √ P        80 MHz to 800 MHz

        E1
d = [ 7 ] √ P         800 MHz to 2.5 GHz                   ……E1
Where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer and d is the recommended separation distance in metres (m).

Field strengths from fixed RF transmitters, as determined by an electromagnetic site survey, a, should be less than the compliance level is each frequency range. b

Interference may occur in the vicinity of equipment marketed with the following symbol:

[image: image49.jpg])





	NOTE 1 At 80 MHz and 800 MHz, the higher frequency range applies.  

NOTE 2 These guidelines may not apply in all situations.  Electromagnetic propagation is affected by absorption and reflection from structures, objects and people.

	a     Field strengths from fixed transmitters, such as base stations for radio (cellular / cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast and TV broadcast cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RD transmitters, an electromagnetic site survey should be considered.  If the measured field strength in the location in which the CO Check is used exceeds the applicable RD compliance level above, the CO Check should be observed to verify normal operation.  If abnormal performance is observed, additional measures may be necessary, such as re-orienting or relocating the CO Check. 

 b     Over the frequency range 150 KHz to 80 MHz, field strength should be less that [V1] V/m


Appendix A – CO Check and Smoking Cessation

The CO Check Pro (Cat No. CO20) / CO Check Baby (Cat No. CO30) are innovative Carbon Monoxide monitors designed specifically for use in smoking cessation programmes and clinics.

All patients attending a smoking cessation programme should have their smoking habits established and checked with the CO Check Pro at every visit.  The results of an initial test may sound alarming to many smokers but within a few days of stopping, CO levels can drop to normal and it is very encouraging for a smoker to see this.  CO monitoring is an extremely powerful tool for a smoking cessation counsellor, GP or respiratory physician.

Taking a test before a COPD patient stops smoking will help you gauge their level of nicotine dependence, and taking it after they have stopped will give them real proof that what they are doing is working.

The best time to do a test is in the afternoon as CO levels fall over-night and morning readings can sometimes give misleadingly low results.  High readings in the morning, on the other hand, would be strong evidence of heavy inhalation and high nicotine dependence.

One single breath of expiration into the CO Check Pro will give instant results in PPM, %COHB and %FCOHb.  Built-in colour light indicators will also give an instant visual display.
Appendix B – CO Check and Environmental Capture  

When the Environmental reading is enabled, background CO levels are taken into account when carrying out a breath test. If the background levels are 10ppm for example, then when the person is taking a breath test this background reading of 10ppm will be taken into account.
For example, if the Environmental monitor is ON and a reading of 10ppm is detected and the patient records 3ppm for example, then the measurement displayed will be the higher value i.e. 10ppm.

Conversely, if the Environmental monitor is OFF, then the measurement displayed will be value recorded by the patient i.e. 3ppm.
[image: image50.png]



19
19

