VM3COP61.11 Production Procedure for Microstim DB3 Tester (2510005)

| Parts list |  |  |
| --- | --- | --- |
| Qty | Description | Part No. |
| 1 | Key fob case | 0131352 |
| 1 | PCB |  |
| 1 | BNC socket |  |
| 1 | LED |  |
| 1 | LED surround |  |
| 1 | Microstim lead | 2520000 |
| 1 | Strain relief | 0010618 |
| 1 | R1 |  |
| 1 | R2 |  |
| 1 | R3 |  |
| 1 | R4 |  |
| 1 | TH1 |  |
| 1 | D1 |  |
| 10mm | Heat shrink (6.0mm) | 0032321 |
| 1 | Cable clamp | 0030500 |

Tool list

Soldering iron

Wire cutters

Stanley knife

Snipe nose pliers

Phillips screw driver (size 0)

Low melt glue gun

Reaming tool

Cable clamp pliers

| - Take the key-fob case and mark a pilot hole as shown. |  |
| --- | --- |
| - Using a small handheld drill, drill through the case. |  |
| - Using a reaming tool, enlarge the hole to 10mm to fit the BNC connector. |  |
| - Remove the keyring bar. |  |
| - File smooth any remnants of the keyring bar. |  |
| - Mark a pilot hole for the strain relief. |  |
| - Using a reaming tool, enlarge the hole to 6mm to fit the strain relief. |  |
| - Affix the top label to the case. - In the centre of the label hole mark a pilot hole for the LED. |  |
| - Using a small handheld drill, drill through the case. - Using a reaming tool, enlarge the hole to 8mm to fit the LED surround. |  |
| - Insert the LED surround and screw on the back nut. |  |
| - Using an abrasive PCB cleaner, clean the PCB solder contacts. |  |
| - Populate the PCB with components as shown. |  |
| - Solder the BNC connector to the PCB. |  |
| - Trim the press-studs from the cable. - Trim the cable to 75cm. |  |
| - Insert the strain relief into the hole at the back of the case. - Attach the cable clamp 35mm from the end of the cable. - Strip and tin 5mm on each of the wires. |  |
| - Apply 10mm of 6mm heat shrink to the cable clamp. |  |
| - Insert the LED as shown. |  |
| - Solder the LED in place. The LED should be 7mm above the PCB. |  |
| - Solder the lead to the PCB. The red lead to the positive contact and the black lead to the negative contact. - Insert the PCB into the case. |  |
| - Test the unit is working by attaching a Microstim DB3 and testing its functions. |  |
| - Using low-melt glue, affix the PCB into the case. |  |
| - Attach the case top and insert the retaining screw. - Affix the serial number label to the underside of the case. - The unit is now ready for QA. |  |