	[image: image1.jpg]



	[image: image8.png]



	Resuscitation Cabinet Installation Instructions

Version 7.0.  11/08/05

Part no. 0391050   


[image: image8.png]Resuscitation Cabinet Installation Instructions

Version 7.0.  11/08/05

Part no. 0391050    


	INFANT RESUSCITATION CABINET

INSTALLATION INSTRUCTIONS




	[image: image10.png]¥




	[image: image2.jpg]





CE0086
	INDEX


Section









Page
1. 
Introduction








2



Important Notice



Overall Installation (diagram)




3



Mounting Heights

2.
Cabinet








4



Installation Notes



Gas Pipe Outlets



Cabinet Dimensions; Viewed from Rear (diagram)

5



Post-Installation Cleaning








3. 
Ceratherm 600-2 Infant Radiant Warmer



6



Electrical Considerations



Weights of Component Parts



Caution!



Radiant Warmer Dimensions in millimetres (diagram)

7



Wall Bracket Mounting Instructions (diagram)


Connecting the Mounting Plate to the Warmer



Mounting the Arm onto the Wall Bracket



8


Hinged Mounting Arm (diagram)








Connecting the Mounting Arm Assembly to the Warmer

4.
Tom Thumb Infant Resuscitation Unit




9

5.
Low Suction Controller






9

6.
Suction Receiving Jar






9

7. 
Digital Apgar Timer







10


8.
Storage Bins








10

9.
Bed & Mattress







10

10.
Commissioning Guidelines






11

11.
Consumables and Parts Listing





13

12.
Warranty








14

13.
Company Details







14






	1.  Introduction


The Viamed Infant Resuscitation Cabinet contains a carefully selected group of products; each main item can be purchased as a single item for use individually. 

Viamed have combined these devices into one product in an arrangement ideally suited to Maternity and Neonatal Intensive Care environments.

Devices forming component parts of the Resuscitation Cabinet: 

· Cabinet with drop down door, housing integral bed and mattress

· Wall-mounted Ceratherm 600-2 radiant warmer

· Tom Thumb infant resuscitation unit

· Low suction controller

· Suction receiving jar

· Digital Apgar timer

· Storage bins

All the components carry the correct and relevant CE marking.

Important Notice

Viamed does not offer an installation service for the wall-mounted resuscitation cabinet system. It is the responsibility of the customer to ensure safe and secure installation in accordance with their own procedures.

It is strongly advised that a survey is undertaken by a qualified structural engineer to determine the appropriate method of mounting, and to ascertain if any wall-strengthening measures or other issues need to be addressed prior to installation.

Servicing personnel must, at all times, be aware of the potential clinical implications of incorrectly serviced equipment.

Overall Installation
[image: image9.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





A. 1325mm; base of the cabinet to the top of the wall bracket.

B. 1245mm; lower cabinet securing bolt to the top wall bracket-securing bolt.

C. 80mm; base of the cabinet to the lower cabinet-securing bolt.

D. 1100mm; maximum distance between the wall and the tip of the radiant warmer.

E. 430mm; base of radiant warmer to the top of the wall bracket.

F. 850mm; base of radiant warmer and the top of the mattress.

G. 45mm; base of the cabinet and the top of the mattress.

Mounting Heights

It is recommended that the cabinet be installed at a height of approximately 1 metre from the ground to the base of the cabinet; this may be varied according to end-user requirements in order to position the cabinet at a convenient working height.

A minimum distance of 800mm must be maintained between the top of the mattress and the underside of the radiant warmer head-unit. 

It is not recommended that this distance be greater than 900mm, as such a distance may position the warmer beyond reach of the operator.

	2.  Cabinet


The cabinet has been custom designed to Viamed's specifications.

When loaded with standard equipment, the cabinet weighs approximately 25Kg. 

Allowance must be made for the additional weight of the infant, and potential forces applied by staff whilst using the equipment. In use, the cabinet is designed to handle loadings of an additional 25Kg.

Installation Notes

Due to different wall constructions, Viamed do not supply fixings. A qualified installer or structural engineer must determine the appropriate fixings based on the results of a site survey.

The positions of mounting holes detailed in this manual are approximate and may vary due to manufacturing and design changes. Viamed recommends that the installer take the necessary measurements from the cabinet itself before finalising position.

When deciding a position for the cabinet on the wall, consideration should be given to the amount of space required to allow easy access to the patient when in use. The cabinet is best mounted away from side walls, with enough space for a person to stand on either side of the drop-down platform.

Gas Pipe Outlets

The cabinet has 3 gas pipe outlets, located in the midpoint of the cabinet base. In a standard configuration only 2 of these outlets are used, the third is incorporated in the event that it may be required in the future.

The cabinet system should be mounted in close proximity to the wall oxygen and vacuum outlets. The oxygen and suction hoses are 3 metres long, which after exiting the cabinet, leaves a length of approximately 2 metres of each hose external to the cabinet.

Cabinet Dimensions; Viewed from Rear

[image: image3.png]



A. 665mm.              

B. 30mm.

C. 455mm.

D. 80mm.

E. 30mm.
F. 555mm.
G. 615mm.
H.
650mm; minimum width required to accommodate the cabinet. Ideally space should be left on both sides to create a working area.

Post-Installation Cleaning

After the cabinet has been installed, care should be taken to clean the cabinet to ensure no dirt or swarf from the installation process remains in, or around, the cabinet.

	3.  Ceratherm 600-2 Infant Radiant Warmer


Comprehensive manuals covering Operation and Servicing of the Ceratherm 600-2 Infant Radiant Warmer are supplied with the Resuscitation Cabinet.  Please refer to these manuals before use.

Electrical Considerations

Each Ceratherm 600-2 Infant Radiant Warmer is safety tested to Class 1, type B and a test report is supplied. 

It is recommended that the equipment be retested for safety and function after installation, and at least annually thereafter.

The Infant Radiant Warmer hinged arm is supplied with a mains lead with a moulded, 13amp fused mains plug.  The mains lead can be shortened and a replacement plug fitted if required, provided this complies with any regulations in place within the customer’s organisation.  Alternatively, the plug can be removed and the mains lead shortened and connected to a 13amp fused spur type mains outlet.

Weights of Component Parts

The following figures are provided to allow calculations of the loadings. It should be noted that the centre of gravity of the head unit of the radiant warmer is not central, but towards the rear of the unit.

· The hinged mounting arm (excluding wall bracket) weighs 3.3Kg.

· The wall bracket weighs 1.7Kg.

· The radiant warmer head unit weighs 4.85Kg. 

· The distance between the mounting points of the bracket is 0.19m.

Caution!

Handle the warmer with care; the ceramic heating element is fragile. 

Do not touch the heating element as this can adversely affect the performance of the element.

The heating element gets very hot in use; take care if the warmer has recently been used or tested. Use only the side handles to adjust the position of the warmer. 

Radiant Warmer Dimensions in millimetres
[image: image4.jpg]300

ok

/

S

~ 580

150

360

09

90

85

1550




	Wall Bracket Mounting Instructions

Remove the plastic housing from the wall bracket.
Mark the position of the mounting holes in a vertical line.

Secure the bracket to the wall using whatever means most suited to the wall construction.

If the wall is uneven place the spacer plate between the wall and the bracket.

Replace the plastic housing.

The bearing may have a protective plastic collar, which must be removed before placement of the mounting arm onto the bearing.

	
	[image: image5.png]





Connecting the Mounting Plate to the Warmer
Using the 4 hex-bolts provided, connect the square mounting plate to the top of the radiant warmer so that the mains cable faces the rear of the warmer. The mains plug will now be in a position to be connected to the mains input socket of the warmer when required.

Mounting the Arm onto the Wall Bracket

Remove the protective sheath from the cylindrical bearing of the wall bracket; this sheath can be disposed of.

Carefully place the mounting arm onto the bearing and slide it down until it sits firmly on the wall bracket. 

Note: there is no additional securing mechanism; the weight of the warmer and arm assembly is sufficient to provide safe, secure mounting.

	Hinged Mounting Arm

	
	
	

	[image: image6.png]" Plastic ring




	
	[image: image7.png]Top View






Connecting the Mounting Arm Assembly to the Warmer

Remove the retaining screw set in the plastic ring on the mounting arm, taking care not to lose the accompanying washer.

Insert a screwdriver between the metal piece and the plastic tongue, loosen the tongue and push the plastic ring upwards.

Remove the semi-circular metal piece; this allows the protective cap to be removed.
Carefully position the warmer under the mounting arm, and lift until the cylindrical section of the mounting plate on the radiant warmer slides into the socket of the arm.

Insert the semi-circular metal piece; care should be taken to ensure that it fits into the groove in the cylindrical section of the mounting plate, thus holding the warmer head unit onto the mounting arm.

Slide the plastic ring down; when it is in place it will secure the semi-circular metal piece. Insert the retaining screw with washer and tighten.

CAUTION: Continue to support the warmer until the retaining screw is secured, failure to do so may result in damage to the warmer and personal injury. 

	4.   Tom Thumb Infant Resuscitation Unit


The Tom Thumb (model TT490-15 3M) Infant Resuscitation Unit has a 3m high-pressure oxygen hose with integral flowmeter, and requires a standard oxygen wall outlet or oxygen cylinder. 

The cabinet should be sited close enough to the oxygen supply to allow connection of the Tom Thumb. 

The Tom Thumb is supplied with its own universal mounting clamp to allow it to be mounted onto one of the rails within the cabinet.

The Tom Thumb is also available in a configuration, which has a low-pressure ‘fir-tree’ inlet and has no flowmeter (model TT480).  The TT480 requires connecting to an existing flowmeter by using 10mm low-pressure tubing, which needs to be cut to length during the commissioning of the cabinet system. 

	5.  Low Suction Controller


The low suction controller requires a standard wall vacuum supply. The cabinet should be sited close enough to the vacuum supply to allow connection of the low suction controller.

A rail clamp is attached to the rear of the low suction controller to allow it to be mounted onto one of the rails within the cabinet.

	6.  Suction Receiving Jar


The suction receiving jar is a reusable 1 Litre canister, containing a 1 Litre disposable liner, manufactured by Vacsax Ltd. The low suction controller connects to the suction receiving jar using 10mm tubing. 

The reusable canister has an integral V-mount adapter, which sits in a universal rail clamp with V-socket to allow it to be mounted onto one of the rails within the cabinet.

Optional reusable jars are available; these are supplied with a mounting cage with an attached rail clamp (see section 11 for part numbers).

	7.  Digital Apgar Timer


When supplied as part of a resuscitation cabinet system, the digital Apgar timer is supplied complete with a mounting clamp assembly to allow it to be mounted onto one of the rails within the cabinet. If ordering an Apgar timer individually, the mounting clamp assembly must be ordered separately (see section 11 for part numbers).

The timer operates from 4x AA/LR6/MN1500 1.5V alkaline batteries (supplied). To access the battery compartment, the mounting bracket assembly must first be removed.

It should be noted that there is no On/Off switch on this device; instead it enters Stand-by Mode after 60 minutes of operation.

To test the functionality of the Apgar timer: refer to the Operator’s manual.

	8.   Storage Bins


Additional storage is provided by 3 storage bins, which can be hooked onto any of the fixed louvres on the rear wall of the cabinet. Additional bins are available (see section 11 for part number).

	9.   Bed & Mattress


Fixed to the inside of the resuscitation cabinet door is a bed, which provides a horizontal platform when lowered.

The bed is designed with sides that angle outwards slightly, providing good access whilst also ensuring the safety of the patient. 

The front edge of the bed has a lip, which angles up at 90 degrees to prevent the mattress from sliding forwards. The mattress sits on the bed with no need for further securing.

	10.  Commissioning Guidelines


The following instructions are for the installation of the equipment to be contained within the cabinet. It is assumed that the cabinet and warmer are securely mounted at an appropriate height as per the instructions in this manual, and that all structural considerations have been taken into account, such as the strength of the wall, and that adequate fixings have been used.

i. Open the cabinet.

ii. Ensure that a bed is mounted on the drop-down door.

iii. Ensure 3 lengths of rail are mounted to the back of the cabinet on the pre-drilled rack to the left-hand side of the cabinet’s rear wall. These should be spaced so that a section of rail is mounted on the uppermost and lowest available positions, and the third rail midway between the 2.

iv. Depending on the model of low suction controller, the vacuum hose may exit from the back of the device vertically upwards or downwards. For upward hoses: feed the probe end of the (yellow) vacuum hose down behind all 3 rails, centrally between the fixing points. For downward hoses: feed the probe end of the hose behind the middle and lower rails.

v. Remove the grommet from the left-hand hole in the cabinet base, feed the hose through the hole until the suction controller can be mounted onto the left-hand section of the top rail. Ensure that the suction controller is mounted in such a way that the action of closing the door will not cause the side of the bed to make contact with the suction unit, and then secure the rail clamp.

vi. Feed the probe end of the vacuum hose through the grommet. Tip: if this action is difficult, rubbing the grommet between the palms of your hands or repeatedly squashing it will soften the rubber, making the task easier.

vii. Slide the grommet along the hose and fit it into the hole in the cabinet base, this grommet prevents damage to the hose. Insert the probe into the wall vacuum supply and test the suction controller by turning it on.

viii. Feed the probe end of the (white) Tom Thumb oxygen hose down behind all 3 rails between the fixing points.

ix. Remove the grommet from the right-hand hole in the cabinet base, feed the hose through the hole until the Tom Thumb can be mounted onto the right-hand section of the middle rail. Ensure that the Tom Thumb is mounted in such a way that the action of closing the door will not cause the side of the bed to make contact with the Tom Thumb, and then secure the rail clamp.

x. Feed the probe end of the oxygen hose through the grommet.

xi. Slide the grommet along the hose and fit it into the hole in the cabinet base, this grommet prevents damage to the hose. Insert the probe into the oxygen supply and test the Tom Thumb by opening the flowmeter. Note: further test and calibration procedures for the Tom Thumb are detailed in separate documentation.

xii. Attach the rail-mounted V-clamp to the bottom rail, on the left hand side.

xiii. If using the Vacsax disposable system: place the tapered connector into the disposable liner in the hole on the top labelled ‘vacuum’. Place the liner into the canister, and mount the canister on the V-clamp on the bottom rail beneath the suction controller. (Note: if using a reusable jar system this will most likely need to be mounted directly onto the middle rail using a jar cage with an integral rail clamp, and the V-mount will not be supplied).

xiv. Cut a length of the clear 10mm suction tubing so as to allow connection of the suction controller to the taper on the disposable liner without excessive coiling or being restrictively short, and connect these devices together.

xv. Connect the long remaining section of the 10mm tubing (approximately 90cm long) to the connection on the disposable liner marked ‘patient’.

xvi. If the digital timer is not operating, remove the mounting bracket and open the battery door, check batteries are correctly inserted. (Note: an Allen Key will be required to perform this action). Place the timer on the right-hand side of the bottom rail, and test the timer functions of Start, Freeze and Reset. (Note: the timer cannot be turned off, instead it enters standby mode after 60 minutes).

xvii. Place the storage bins onto the louvred rails on the right-hand side of the rear wall of the cabinet.

xviii. Place the mattress on the bed; if the mattress is wrapped in a plastic bag or film, ensure that this is removed. 

xix. Place a Tom Thumb patient circuit into the cabinet; the shelf is the ideal place for it. The bag need not be opened until the system is being prepared for use.

xx. Place the Operator’s manuals for the resuscitation cabinet and radiant warmer in the bottom of the cabinet beneath the storage bins.

xxi. Ensure the cabinet closes easily, without obstruction or trapping anything in the hinge of the door as it closes. Note: if any pressure is required to hold the door closed in order to secure the hinge, it is likely that something is trapped, and forcing the cabinet closed in this way can lead to damage to the hinge. 

xxii. If it is felt that the trailing hoses are unsightly or causing an obstruction, use tie-wraps to create loops as close to the base of the cabinet as possible, if these loops are of the right size it is possible to conceal and secure them using the flange underneath the base of the cabinet to hold them in place.

	11.  Consumables and Parts Listing


Note: the Viamed resuscitation cabinet is available in different configurations, the parts list contains all variants; care must be taken when re-ordering replacement parts. If in doubt, please contact Viamed to determine the correct item.

	Consumables

	Part number
	Description

	0350040
	Tom Thumb patient circuits.

Includes face mask.

Single patient use.

	9910-340
	Vacsax disposable liners.

1 Litre, standard bore. Single patient use.

Box of 25.

	Parts

	Part number
	Description

	0310030
	Tom Thumb infant resuscitator, model TT480.

No flowmeter.

	0310034
	Tom Thumb infant resuscitator, model TT490-15 3M.

Incorporates flowmeter, with 3m high-pressure hose and probe for connection to wall oxygen outlet. 

	0310037
	Therapy Equipment low suction controller.

Includes 3m vacuum hose with rail clamp.

	0310035
	Oxylitre low suction controller.

(Note: Vacuum hose, 0330040, supplied separately).

	0330040
	Oxylitre vacuum hose, 3m, includes rail clamp.

For use with Oxylitre low suction controller.

	3383-132
	Vacsax suction receiving canister.

1 Litre, standard bore. Reusable. With dovetail mount.

(Note: mounting bracket supplied separately)

	0320010
	Universal rail clamp with ‘V’.

For mounting Vacsax suction jar.

	3833-007
	Vacsax suction tapers.

Connects Vacsax disposable liner to low suction controller via 10mm tubing.

Pack of 10.

	0320015
	Reusable ½ Litre suction jar. Requires holder, 0320025.

	0320025
	Holder for ½ Litre suction jar, including rail clamp, for use with 0320015 jar.

	0310100
	Digital Apgar timer.

Features large-digit display, time-freeze function, and audible indication at 1, 5 and 10 minutes elapsed time.

(Note: mounting bracket supplied separately)

	0320200
	Mounting bracket with rail clamp.

For mounting digital Apgar timer.

	0320021
	Resuscitation bed, 35cm (14”) wide.

With sides and front retaining lip to prevent mattress movement.

	0320043
	Mattress, 35cm (14”) wide. 

Designed to meet BS5852 (Crib5) fire retardancy standards.

For use with 0320021: 35cm wide bed.

	0310020
	Storage bin, blue.

	0330120
	Radiant warmer element, 600W.


	12.  Warranty


All parts have a warranty period of 12 months, beginning from the date of invoice. 

This warranty covers defects in materials and workmanship, but excludes accidental damage and damage caused by misuse or neglect. 

Should any component develop a defect within the warranty period it will, at Viamed’s discretion, be repaired or replaced.

For products purchased outside the UK, please contact your local distributor.

	13.  Company Details


All products are CE marked to the requirements of MDD 93/42 EEC and are supplied in accordance with our quality system accreditations:
BS EN ISO 9001/2000, ISO 13485/2003 and CMDCAS.

Contact details:

Viamed Ltd,

15 Station Road,

Cross Hills,

Keighley,

West Yorkshire,

BD20 7DT.

United Kingdom.

Tel:  +44 (0)1535 634542

Fax: +44 (0)1535 635582

Email: info@viamed.co.uk

Web: www.viamed.co.uk



































































© Copyright 2005. Viamed Ltd.

1

