VM3COP40.14 R-22DEM Oxygen Sensor Production Procedure

| Parts list |  |  |
| --- | --- | --- |
| Qty | Description | Part No. |
| 1 | Oxygen sensor – R-22AV | 0110132 |
| 1 | Connector Housing, M20 Series, 2 Way | 9081047 |
| 2 | M20 Series Female Crimp Contacts | 9071046 |
| 1 | Anti-static gas barrier bag | 0150000 |
| 1 | R-22DEM PCB | 9015004 |
| 2 | Resistor 4k7Ω surface mount | - |
| 5cm | Red wire | - |
| 5cm | Black wire | - |

Tool list

Soldering iron

Wire cutters

Stanley knife

Small pliers

Low melt glue gun

Abrasive contact cleaner

M20 crimp tool

Wire stripper

Strip heat sealer

| PCB Preparation: |  |
| --- | --- |
| - Using an abrasive contact cleaner, clean the solder contact surfaces of the PCB. |  |
| - Solder the two resistors into place as shown. Note: The resistor are surface mount components, it may be necessary to use magnification when soldering or inspecting the PCB. |  |

| Sensor production: |  |
| --- | --- |
| - Remove the sensor from the sealed protective bag. |  |
| - Remove the top cap from the sensor. Use VM3COP40.11 for the method. |  |
| - Remove the white plastic surround from the Molex connector. |  |
| - Noting the Red and Black marking on the sensor surface, place the PCB over the sensor’s pin contacts as shown. |  |
| - Solder the pins to PCB contacts. - Trim the sensor’s pins as shown. |  |
| - Strip 3mm of cable outer from each end of the black and red wires. - Using the M20 crimp tool, affix a pin to one end of red and black wires. Tin the remaining end of each wire. - Insert the pins into the connector housing as shown. |  |
| - Noting the Red and Black marking on the sensor surface, solder the connector assembly to the PCB as shown. - Test the sensor by connecting it to a digital volt meter. The output in air should be between 4.0mV and 6.5mV. |  |
| - Using a low melt glue gun, cover the small PCB and its components in adhesive and allow to set. |  |
| - Noting the serial number of the sensor, remove the sensor label. Replace the label with the Vandagraph label with the corresponding serial number. |  |
| - The sensor is now ready to be QA tested. As before, test the sensor by connecting it to a digital volt meter. The output in air should be between 4.0mV and 6.5mV. - Record the QA results on the Intrastats system. |  |
| - Place the sensor into a gas barrier bag. - Seal the bag using a strip heat sealer. - Affix the corresponding serial number to the bag. |  |