VM3COP51.30 Resus Cabinet Assembly

| Parts list |  |  |
| --- | --- | --- |
| Qty | Description | Part No. |
| 1 | Resuscitation Cabinet | 0330105 |
| 320mm | Edging Grommet Strip 1.6mm to 2.0mm | 0330012 |
| 1 | Resuscitation Bed 14” flat | 0320021 |
| 6 | Self-Tapping Screw No. 6 x 0.5 inch | 0330064 |
| 1 | Mattress large - infant resuscitation | 0320043 |
| 1 | Resuscitation Cabinet Rail - 0.3m length | 0332058 |
| 2 | Resuscitation Cabinet Rail - 0.4m length | 0332059 |
| 8 | Resuscitation Cabinet Rail - Spacer 20mm | 0332039 |
| 8 | Resuscitation Cabinet Rail - Bolt 35mm | 0332041 |
| 8 | Resuscitation Cabinet Rail - Lock nut | 0332043 |
| 3 | Storage bins: Size 3 (blue) | 0310020 |
| 1 | Label: Viamed (domed) | 0390056 |
| 1 | Label: Support door (domed) | 0390057 |
| 1 | Label: Caution | 0390054 |
| 1 | Label: Serial number | 0390053 |

Tool list

Pozi-drive No.2

Pozi-drive No.3

Drill with 2mm bit

10mm open ended spanner

Ruler

Pencil

Hand held vacuum cleaner

| - Unpack and inspect the cabinet for damage. Faults or damage must be reported to the warehouse team leader. |  |
| --- | --- |
| - Affix the front labels as shown. These must be central and straight. |  |
| - Affix the serial number label to the right-hand side of the cabinet as shown. This should be central and 90mm from the base of the cabinet. |  |
| - Affix the rubber door stoppers to either side of the door opening. |  |
| - Cut a length of edging strip to 320mm. - Insert the protective strip around the edges of the inlet access hole located in the base of the cabinet. |  |
| - Inside the cabinet, on the rear panel, are three rail mounts. Place a nut behind the rail mount. |  |
| - Place piece of folded card behind the nut to hold it in place. The standard arrangement for these is: Top, bottom and fifth from bottom. The rails are of two lengths. The top one is short, the two lower rails are long. |  |
| - Place a spacer on the top side of the rail mount. Ensure that the spacer aligns with the nut. - Repeat this process on each rail. |  |
| - Place a counter sunk bolt through each hole in the rail. Note: The rail has channel on the lower side. |  |
| - Affix each rail to the rail mounts. The rails can be fitted in several positions to suit customer requirements. The standard positions are: Top – short rail One below middle – Long rail Bottom – Long rail |  |
| - Using a pencil, mark 50mm from the left side of the cabinet door. This measurement must be taken from the panel and not the rubber edge strip. |  |
| - Place the bed on the cabinet door. And align with the pencil mark. |  |
| - Align the top of the bed with the top of the door panel. This alignment must be from the panel and not the rubber edge strip. |  |
| - Using a pencil, mark through each of the six holes in the bottom of the bed. |  |
| - Using a 2mm drill bit, drill pilot holes to a depth of 10mm at each mark. - Using a hand-held vacuum cleaner, remove any swarf. |  |
| - Using the six fixing screws, affix the bed in place. - Ensure that the door closes with the bed in place. |  |
| - Leave the mattress in the original packaging and place on top of the bed. |  |
| - Wrap the three plastic storage bins in bubble wrap and place them on the shelf in the top right corner. |  |
| - Affix the ‘Caution’ label on the inside of the door. The label should be 5mm from the edges. |  |