| Description. | Tom Thumb. | Time & Date of Calibration / Test. |  |
| --- | --- | --- | --- |
| Model. |  | Time & Date of QA. |  |
| Serial No. |  | Serial no. : O2 flowmeter. |  |
|  |  | Serial no. : Air flowmeter. |  |

Do not start QA check within 1 hour of the calibration / test.  Labelling is to be attached after calibration / test.

Record the manometer reading in millibars below, ensuring the Tom Thumb meets the limits specified.

| Test Equipment | Test | Specification | Reading |  | P/F |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Cal | Q.A. |  |
| Snoop liquid. | Check all ports / connections for leaks. | No bubbling. |  |  |  |
| CE 078. | Adjustable Valve : @ 10 Lpm | Minimum : <8 cmH2O. |  |  |  |
| CE 078. | Adjustable Valve : @ 10 Lpm. | Maximum : 43±1 cmH2O. |  |  |  |
| CE 078. | Precision Valve : @ 10 Lpm. | Maximum : 45±1 cmH2O |  |  |  |
| CE 078. | Pressure Gauge Test. @30cmH2O | Pressure : 30cmH2O ±0.5 |  |  |  |
| Visual check. | Gauge cannot be removed without tools. |  |  |  |  |
| Visual check. | Gauge appears straight. |  |  |  |  |
| Visual check. | All adjustable settings are set to a minimum. |  |  |  |  |
| Visual check. | Labels are attached. | CE label. |  |  |  |
|  |  | Serial No. label. |  |  |  |
|  |  | Tom Thumb label |  |  |  |

Calibration :					Q.A check :

Signed :  …………………………………                   Signed :  ………………………………