VM3COP43.07 Maxtec Ultramax Analyser Testing

| - Unbox the Ultramax analyser. |  |
| --- | --- |
| - Check the box contains all the items listed: Ultramax analyser Operating manual Clear tubing Soft cover Batteries |  |
| - Using a #1 Phillips screw driver, remove the battery door and insert two AA batteries. Battery orientation markings can be found on the inside of the battery compartment. |  |
| - Activate the analyser by pressing the power button on the front panel. |  |
| - Connect one end the tubing to gas inlet port. |  |
| - Connect the remaining end of the tubing to gas source with an inline flow meter. |  |
| - Test the flow meter by allowing gas to flow at the rate of 2 litres per minute (lpm). Compare the reading on the Ultramax monitor to the reading on the inline flow meter. A reading of ±0.2lpm is acceptable. - Repeat the process at 5 and 8 lpm. |  |
| - Test the PSI gas pressure function. Allow gas to flow at 2lpm. - Cover the gas sample outlet port, located on the right-hand side, with a finger or thumb. - Allow 5 seconds before observing a pressure reading. |  |
| - Set the pressure mode switch kPa. This can be found inside the battery compartment. |  |
| - Test the kPa gas pressure function. Allow gas to flow at 2lpm - Cover the gas sample outlet port, located on the right-hand side, with a finger or thumb. - Allow 5 seconds before observing a pressure reading. - Return the mode switch to the PSI setting. |  |
| - Verify the Oxygen calibration. Connect the inlet tubing to a source of pure Oxygen (≥99.95%). Allow gas to flow at 2lpm. - Press and hold the mode button. While holding the mode button, the gas measurement should read between 98.5 and 101.5%. If the gas measurement is not within this range, then inform the team leader. - Record the QA results on the intrastats system. |  |