CLINICAL EVALUATION REPORT
Compiled by
Derek Lamb
Microstim MkIII

# 1. General details
   Microstim MkIII Peripheral nerve stimulator designed for use during anaesthesia.
 Manufacturer:
Viamed Ltd
15 Station Road
Cross Hills
Keighley
West Yorkshire
BD20 7DT

# 2. Description of the device and its intended application.Intended therapeutic and/or diagnostic indications and claims

Muscle relaxants, often used with general anaesthetics during surgery, block neuromuscular
transmission (the communication process whereby a nerve impulse produces muscle
contractions), and therefore the sensation of feeling.
The Microstim is a peripheral nerve stimulator used to monitor this block by simulating a stimulus
from the brain to the muscle.
Any resultant muscle contraction indicates the patient’s level of paralysis, the anaesthetist can
then evaluate the neuromuscular block (and therefore the sensation of feeling), and take
appropriate action.
# 4. Context of the evaluation and choice of clinical data types.
The Clinic Evaluation has been performed via, Clinical Experience data / documentation and Clinical literature search(s).

Sources of data/documentation used in clinical evaluation.

Review of Essential requirements Current DOCID 15722
Review Risk Assessment current DOCID 19657
Risk against Benefits DOCID 16895
Post-market Surveillance report:
Section 1 Stock Identification
Section 2 Supplier Review
Section 3 Sales Review
Section 4 Countries Review
Section 5 Returns / Services Review
Section 6 Design Changes Review
Section 7 IFU Review
Section 8 Labels Review
Section 9 Documentation updates Review
Section 10 Internal Issues Review
Section 11 Clinical Data / FDA Incidents Search
Current DOCID 19655

All returns / failure records since 2001 are computerized and reviewed annually.
See post market surveillance reports DOCID 19655.
There have been no recalls of Microstim units.
All units have Q.A. Records and service records indicating the devices perform as intended.
Sales / Returns and service logs data See DOCID 19655
Microstim units have been on the Market since 1990's. There have been no signification	design changes in 10-20 years..
# 5. Summary of the clinical data and appraisal

As there have been no clinical trials carried out on behalf of Viamed, this report is based on literature search . Technical File “Microstim“
As Per Annex X CLINICAL EVALUATION
1.1d Where demonstration of conformity with essential requirements based on clinical data is not deemed appropriate, adequate justification for any such exclusion has to be given based on risk management output and under consideration of the specifics of the device/body interaction, the clinical performances intended and the claims of the manufacturer. Adequacy of demonstration of conformity with the essential requirements by performance evaluation, bench testing and pre-clinical evaluation alone has to be duly substantiated.