| VM3COP40.20 R-33VAN-2 Production Procedure |  |  |  |  |
| --- | --- | --- | --- | --- |
| Parts list: |  |  |  |  |
| Description | Part Number | Supplier | Supplier part number | Quantity |
| R-33S Sensor | 0110226 | Teledyne | R-33S | 1 |
| Black housing – R-33VAN | 9720200 | Vandagraph |  | 1 |
| TDK ferrite - 2035-0930 |  | Farnell | 1301656 | 1 |
| Cable | 97300500 | VDC | 211-311-000 | 1m |
| Terminal pin - Multicomp E0508 | 97300503 | Farnell | 9972218 | 2 |
| Viton ‘O’ ring (Large) | 97300501 | Spen Bearings | BS0140-15 | 1 |
| Viton ‘O’ ring (Small) |  | Spen Bearings | VIT-0130:10-274 |  |
| Sensor top cap |  | Vandagraph |  | 1 |
| Grommet | 9730015 |  |  | 1 |
| Cable grip |  |  |  | 1 |

Tool list

Soldering iron

Wire cutters

Wire stripper

Stanley knife

Heat gun

Crimp tool

Scissors

| - Remove the R-33S Oxygen sensor from the packaging. |  |
| --- | --- |
| - De-solder the red and black wires. - Remove the label from the sensor. Note the serial number. When the sensor is relabelled, the same serial number will need to be applied. |  |
| - Insert a grommet into the hole in top cap. |  |
| - Feed the cable through the grommet. |  |
| - Strip 12mm from the outer cable sheath. - Cut the shield away. - Affix a cable grip 2mm from the cable end. |  |
| - Place 10mm of 6mm heat shrink over the cable grip. Using a heat gun, shrink into place. |  |
| - Trim the sensor housing to leave 18mm on the wider part. |  |
| - Apply clear adhesive silicone to the inside face of the sensor housing. |  |
| - Place the sensor inside the sensor housing and firmly press down. |  |
| - Solder the red and black wires to the corresponding contacts on the PCB. |  |
| - Apply a modest amount of super glue to the inside of the sensor cap. Firmly press and hold the sensor cap in place. Wipe away any excess. |  |
| - Affix the serial number label. |  |
| - Place the large ‘O’ ring over gas inlet port. - Place the small ‘O’ ring in the recess on the length of the gas inlet port. |  |
| - Strip 70mm form the outer sheath of the cable and cut away the shield. - Strip and tin 7mm from each wire - Using a crimp tool, affix a terminal pin to each wire. |  |
| - Cut 20mm of 6mm heat-shrink. Using a heat gun, shrink into place. |  |
| - Affix a ferrite by looping once thought. Leave 60mm of sheathed cable protruding from the ferrite. |  |
| Q.A. |  |
| - Using a digital volt meter, test the output of the sensor. The output must be 25mV ±2mV. - Note the output and record the results on the Intrastats system. |  |
| Packing |  |
| - Place the sensor in a gas barrier bag. - Using a strip heat sealer, seal the sensor within the bag. - Affix the corresponding serial number label to the outside of the bag. |  |