Device verification

Test Report ISO 80601-2-55
DOCID 20842

Derek Lamb
01/08/2017

The ISO 60601-1 would be applied for complete medical electrical devices or systems. The oxygen sensor is only a component of medical electrical devices. That’s also the reason for not applicable EMC test only for the sensor.
So we have the certificate of conformance according to ISO 80601-2-55 in reference to specific particular requirements for the basic safety and essential performance of a respiratory gas monitor (RGM) and here the part of the oxygen measurement related to the oxygen sensor.

Regards,
Marko

ppa. Marko Sprössel
R&D Manager
Honeywell I Safety & Productivity Solutions
Intelligent Life Care
EnviteC | Ackermann
Office:  +49 3841 360 313
Mobile: +49 176 12343213