VM3COP37.56 0110525 KE-25 Oxygen sensor QA Procedure

| Tools required: |  |
| --- | --- |
|  | Description: |
|  | Multi-meter |
|  | KE-25 multi-meter adapter |

| Method: |  |
| --- | --- |
| - Connect the KE-25 multi-meter adapter to the multi-meter, ensuring that the polarity is correct (Black to common) |  |
| - Set the multi-meter to the millivolt range. |  |
| - Connect to KE-25 sensor to the multi-meter adapter. The output reading, to pass QA, should be between 10.0mV and 15.5mV. |  |
| - Fill out the QA sheet that came with the production job. Fill out the QA form on intrastats. The sensor is now ready for packaging. |  |