VM3COP37.58 0110436 MAX-250B Oxygen Sensor QA Procedure

| Tools required: |  |
| --- | --- |
|  | Description: |
|  | Multi-meter |
|  | MAX-250B multi-meter adapter |

| - Connect the MAX-250B multi-meter adapter to the multi-meter, ensuring that the polarity is correct (Black to common) |  |
| --- | --- |
| - Set the multi-meter to the millivolt range. |  |
| - Connect to MAX-250B sensor to the multi-meter adapter. The output reading, to pass QA, should be between 10.0mV and 15.5mV. |  |
| - Fill out the QA sheet that came with the production job. Fill out the QA form on intrastats. The sensor is now ready for packaging. |  |