VM3COP37.60 0110437 Maxtec Oxygen Sensor Production Procedure

| Tools required: |  |  |
| --- | --- | --- |
|  |  | Description: |
|  |  | Soldering iron Long nose pliers Small vice Wire cutters |
| Parts list: |  |  |
|  | Part number: | Description: |
|  | 0110432 | Oxygen sensor MAX-250B |
|  | 0131525 | Connector 2 pole black socket 1mm |
|  | 0131526 | Connector hood black |
|  | 0032310 | Heat shrink tubing – black 1.6mm |

| - Remove the sensor from the outer packaging. |  |
| --- | --- |
| - Place the connector hood over both sensor wires. |  |
| - Strip and tin 2mm of each wire. Place 5mm of 1.6mm heat shrink over each wire. |  |
| - Apply solder to the connector pins. |  |
| - Solder the wires to the connector pins. |  |
| - Cover the pins with the heat shrink. Shrink the tubing leaving none of the pin exposed. |  |
| - Screw the connector hood to the connector. If the wires are not compact enough then there may be resistance when tightening the hood, this can cause the pins to bend. |  |
| - Affix a serial number label to the sensor. The sensor is now ready for QA. |  |