Certificate of Calibration

Certificate number:
Tom Thumb Infant Resuscitator.

Part Number: 

Serial number:

This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with Tom Thumb service manual.

	Equipment used:
	Digitron Digital Manometer P200 (CE    )

	
	Digitron Digital Manometer P200 (CE    )

	
	Flow Meter (CE130)

	
	Snoop liquid


	Test Equipment
	Test
	Specification
	Reading
	Pass/Fail

	
	
	
	
	

	Snoop liquid
	Check all ports / connections for leaks
	No bubbling
	N/A
	

	
	
	
	
	

	Digital Manometer (CE***)
	Adjustable Valve:

@ 10 Lpm
	Minimum:

<8 cmH2O
	
	

	
	
	
	
	

	Digital Manometer (CE***)
	Adjustable Valve:

@ 10 Lpm
	Maximum:

43±1 cmH2O
	
	

	
	
	
	
	

	Digital Manometer (CE***)
	Precision Valve:

@ 10 Lpm
	Maximum:

45±1 cmH2O
	
	

	
	
	
	
	

	Digital Manometer (CE***)
	Pressure Gauge Test

@30cmH2O
	Pressure:

30cmH2O ±0.5
	
	


Name:
 
    

Date:


Signed: 
______________________                                                                                         

