[image: image3.jpg]10,9
' I ‘
N
. Sl -
< ™ g
I o ™
Y
M16x1 (5/8-24 UNEF)*
28
@29






Specification Datasheet - 8010034
[image: image4.jpg]Radiall R114.426.000 SMB



Fathom Dive Systems LLC. 

FDS100 Oxygen Sensor
© Vandagraph Sensor Technologies Ltd.

v1.0
17th March 2026
	Specifications


	VST part number
	8010034

	OEM model number
	FDS100 Oxygen Cell

	Initial output signal in ambient air
	9 – 13 mV

	Output connector
	Radiall SMB Series 

R114.426.000 straight male receptacle for PCB, solder legs

	Input
	Threaded M16 x 1mm

	Input O ring
	No

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1 hour equilibration time)

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25 ºC after 36 seconds

	Cross interference
	< 0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	≥ 500,000 % oxygen hours

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Shelf life
	24 months

	Load required
	≥ 10 K Ohms

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes 

	Pressure equalization - PCB
	Yes

	Pressure equalization - output cap
	None

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag - clear

	Warranty period
	12 months from date of sales invoice

	All specifications are applicable at standard conditions: 1013 mbar, 25 °C dry ambient air.



Page 2 of 4

	Drawings/images


[image: image5.png]



[image: image6.png]



[image: image7.png]


[image: image8.png]Vandagraph
>ensor Technologies





Page 3 of 4

	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image1.png]FATHOM ARCE ovonser

S Y ST EMS SIN FDS123456

www fathomdive.com ] 2026-01 (IR T 0





Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.png][FATIHOM

E S Y S T E

www.fathomdive.com

FDS100 Oxygen Sensor

(M| 2026-01

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with
eyes and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor

from this bag until ready to use. Depending upon the ambient conditions the oxygen
sensor may require 30 minutes fo reach signal stability after installation.





 NOTE: Serial number prefix to be  FDS e.g. FDS123456
Page 4 of 4
Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

