Certificate of Calibration

Certificate number: 1505125548
V1000 Foetal Heart Simulator

Part Number: 1410000

Serial number:PR02202A10
This certificate reports recorded values for the instrument specified above.

The output rates were measured in accordance with procedure: VM3COP55.01. 
	Equipment used:
	Sonicade FM6 foetal monitor (CE051)

	
	

	
	TTi TF930 3GHz counter (CE185) - calibration traceable to UKAS calibration standards.


	V1000 

BPM setting


	Period values stated in milliseconds
	Result

PASS/FAIL

	
	Target
	Lower limit
	Upper limit
	Actual
	

	60
	1000.00
	990.00
	1010.00
	992.66
	PASS

	90
	666.70
	660.03
	673.37
	665.33
	PASS

	120
	500.00
	495.00
	505.00
	499.80
	PASS

	150
	400.00
	396.00
	404.00
	399.95
	PASS

	180
	333.30
	329.97
	336.63
	332.47
	PASS

	210
	285.70
	282.84
	288.25
	285.69
	PASS


Name:
 M.Green
Date:
11/09/2017
Signed: 
______________________                                                                                         

