Apgar Timer Production Procedure

VM3COP47.03

Tool list:

Soldering iron

Side cutters

Long nose pliers

Torx T-10 driver

Spanner 5.5mm

| Parts list: | Quantity: | Part number: |
| --- | --- | --- |
| Label top panel | 1 | 0330505 |
| Label bottom panel | 1 | 0330506 |
| Apgar case | 1 | 0330548 |
| Apgar case end | 2 | 0330549 |
| Apgar timer PCB v1.0 | 1 | 0330555 |
| Apgar battery holder | 1 | 0330558 |
| PCB spacer – single unit | 2 | 0330566 |
| Apgar keypad membrane | 1 | 0330570 |
| Apgar timer LCD | 1 | 0330575 |

| Production method: |  |
| --- | --- |
| - Apply the, 1.5V AA, battery orientation labels to the battery box and insert the battery contacts. |  |
| - Insert the battery box into the rear case panel. |  |
| - Peel away the adhesive backing from the keypad. Route the keypad connector through the small hole and stick the keypad down. |  |
| - Note the serial number of the PCB, which can be found on the lower left front side of the PCB. |  |
| - Place the PCB over the battery contacts. |  |
| - Ensure that the PCB is flush to the battery box then solder the contacts to the PCB. |  |
| - Trim away the excess battery contact tab. |  |
| - Locate the display orientation markers on the PCB and LCD. |  |
| - With the orientation markers aligned, insert the LCD contacts into the PCB. |  |
| - Support the display at each end. |  |
| - Solder the LCD to the PCB then remove the supports. |  |
| - Connect the keypad connector to the PCB. At this point, the unit can be tested. Refer to the testing section, on page 6, in this manual. |  |
| - Place the spacers over the retaining screws on the front case. |  |
| - Remove the protective film from the LCD. |  |
| - Guide the screws through the holes in the PCB then snap the two case halves together. |  |
| - Using a 5.5mm spanner, affix the retaining nuts to each screw thread. |  |
| - Position the top cap on the case, ensure that the seal is correctly located. Using a torx T-10 driver, screw the cap down. Repeat the process for the bottom cap. |  |
| - Insert the screw covers to all four screw holes. |  |
| - Affix the top label. |  |
| - Affix the bottom label. |  |