BS EN 60601-2-10:2015

Does not apply to the Supramaximal Peripheral Nerve Stimultors.

Rational:

 Currently there isn`t a specified particular requirements standard for Supramaximal Peripheral Nerve Stimultors.
BS EN 60601-2-10:2015 Particular requirements for the basic safety and essential performance of nerve and muscle stimulators.
This does not apply the Supramaximal Peripheral Nerve Stimultors by virtue of BS EN 60601-2-10:2015 definition that just includes transcutaneous electrical nerve stimulators (TENS) and electrical muscle stimulators (EMS). This is further clarified by subclause 201.3.204:
Stimulator. ME EQUIPMENT for the application of electric currents via electrodes in direct contact with the PATIENT for diagnosis and/or therapy of nuromuscular disorders.
Whereas, Supramaximal peripheral nerve stimulators have been designed to aid anaesthetists to detect and monitor neuromuscular
function. They are used to assess the neuromuscular block during surgery and the reversal of neuromuscular
block during the recovery period. By using a peripheral nerve stimulator a clinician can determine the effects of
neuromuscular-blocking drugs and adjust the dosages accordingly by observing muscular response to
different patterns of electrical nerve stimulation.

See Issue 103724 Steve Nixon 25 / 09 / 2017