SECTION 2 - SAFETY

## 2.1 Intended Use

UT4000A is a multi-parameter patient monitor, providing continuous monitoring, display of electrocardiograph, non-invasive blood pressure, blood oxygen saturation of patients from adults to neonates by means of display on a non-permanent graphical display. This device is for use only by trained medical personnel in the hospitals and ambulances. It is not designed for home use. The device is restricted to be used on one patient at a time.

## 2.2 Terminology

The terms danger, warning, and caution are used throughout this manual to point out hazards and to designate a degree or level of seriousness. Familiarize yourself with their definitions and significance.

Hazard is defined as a source of potential injury to a person.

WARNING indicates a potential hazard or unsafe practice which, if not avoided, could result in death or serious injury.

CAUTION indicates a potential hazard or unsafe practice which, if not avoided, could result in minor personal injury or product /property damage.

NOTE provides application tips or other useful information to assure that you get the most from your equipment.

## 2.3 Monitor Safety

The safety statements presented in this chapter refer to the equipment in general and,

in most cases, apply to all aspects of the monitor. There are additional safety statements in the parameter chapters which are specific to that monitored parameter.

The order in which safety statements are presented in no way implies order of importance.

There are no dangers that refer to the equipment in general. Specific “Danger” statements may be given in the respective sections of this manual.

## 2.4 Safety Requirments

- The following warnings and cautions must be read and understood before operation of the patient monitor.

2.4.1 WARNINGS

- UT 4000A patient monitor is not intended to be used as an apnea monitor.

- UT 4000A patient monitor is not intended to be used during MRI or CT scan.

- OPERATOR – Medical technical equipment such as this monitor must only

be used by persons who have received adequate training in the use of such

equipment and who are capable of applying it properly.

- EXPLOSION HAZARD – Do not use UT4000A patient monitor in a

flammable atmosphere where concentrations of flammable anesthetics or

other materials may occur.

- POWER SUPPLY – The device must be connected to a properly installed

power outlet with protective earth contacts only. If the installation does not

provide earth conductor, disconnect the monitor from the power line and

operate it on battery power, if possible.

- Do not connect to an electrical outlet controlled by a wall switch.

- When using High Frequency surgical equipment, you should put the patient

cables far from the surgical operation in order to reduce the hazards of burns in

the event of a defect in the HF surgical equipment neutral electrode connection.

- Route all cables away from patient’s throat to avoid possible strangulation.

- For safety reasons, all connectors for patient cables and sensor leads are

designed to prevent inadvertent disconnection, should someone pull on them.

Do not route cables in a way that they may present a stumbling hazard. For

device installed above the patient, adequate precautions must be taken to

prevent them from dropping on the patient.

- DEFIBRILLATION – Do not come into contact with patients during

defibrillation, and be sure that the patient cable is not in touch the metal parts.

Otherwise serious injury or death could result.

- Conductive Connections – Extreme care must be exercised when applying

medical electrical equipment. Many parts of the man/machine circuit are

conductive, such as the patient, connectors, electrodes. It is very important that

these conductive parts do not come into contact with other grounded,

conductive parts when connected to the insulated patient input of the device.

Such contact would bridge the patient’s insulation and cancel the protection

provided by the insulated input. In particular, there must be no contact of the

neutral electrode and ground.

- The leakage current will rise up due to simultaneous use of other PATIENT-

connected MEDICAL ELECTRICAL EQUIPMENT, for example, a cardiac

pacemaker or other electrical stimulators. Before connecting each other, the

leakage current test should be done by qualify personnel and the leakage

current should be within the limit of EN 601-1 standard.

- Disconnection From Mains – When disconnecting the system from the power

line, remove the plug from the wall outlet first. Then you may disconnect the

power cord from the device.

- SHOCK HAZARD – Do not attempt to connect or disconnect a power cord

with wet hands. Make certain that your hands are clean and dry before

touching a power cord.

- Electro-surgery return circuit must be connected properly to prevent burns at

monitoring sites and even death.

- Do not touch signal input or output outlets and the patient simultaneously.

- Do not use unprotected leadwires having exposed conductors at the cable end.

The use of unprotected leadwires and cables may pose an unreasonable risk of

adverse health consequences or death.

- Before putting the system into operation visually inspect all connecting cables

for signs of damage. Damaged cables and connectors must be replaced

immediately.

- This patient monitor has no defibrillator synchronization output. Please do not

connect the monitor to a defibrillator.

- ALARMS – Do not rely exclusively on the audible alarm system for patient

monitoring. Ajustment of alarm volume to a low level or off during patient

monitoring may result in a hazard to the patient. Remember that the most

reliable method of patient monitoring combines close personal surveillance

with correct operation of monitoring equipment.

- The functions of the alarm system for monitoring the patient must be verified

at regular intervals.

- Before using the system, the operator must verify that it is in correct working

order and operating condition.

- INTERFACING OTHER EQUIPMENT – Devices may only be interconnected

with each other or to parts of the system when it has been determined by

qualified biomedical engineering personnel that there is no danger to the

patient, the operator, or the environment as a result. In those instances where

there is any element of doubt concerning the safety of connected devices, the

user must contact the manufacturers concerned for proper operation should be

verified with the applicable manufacturer’s instructions for use, and system

standards IEC 60601-1-1/EN 60601-1-1 must be complied with.

- Accessory equipment connected to the analog and digital interfaces must be

certified according to the respective IEC standards ( e.g. IEC 950 for data

processing equipment and IEC 601-1 for medical equipment ). Furthermore,

all configurations shall comply with the valid version of the system standard-

EN 601-1-1. Everybody who connects additional equipment to the signal

input part or signal output part configures a medical system, and is therefore

responsible that the system complies with the requirement of EN 601-1-1.

If in doubt, consult the technical service department or your local

representative.

- In all cases, safe and proper operation should be verified with the applicable

manufacturer's instructions for use and system standards EN 60601-1-1/

EN60601-1-1 must be complied with.

- Every 12 months a calibration of all measuring functions shall be performed by

authorized personnel.

- The monitor is not for home use.

- ACCIDENTAL SPILLS – To avoid electric shock or device malfunction,

liquids must not be allowed to enter the device. If liquids have entered the

device, take it out of service and have it checked by a service technician

before it is used again.

- LEAKAGE CURRENT TEST – When interfacing with other equipment, a test

for leakage current must be performed by qualified biomedical engineering

personnel before using with patients.

- 2.4.2 CAUTIONS:

- EMC – Electromagnetic field may interfere with the proper performance of the

patient monitor. Make sure that all external devices operated in the vicinity of

the monitor comply with the relevant EMC requirements. Also keep cellular

telephones or other telecommunication equipment away from the monitor.

- ECG patient cable shall be connected to the ECG leadwire supplied with this

monitor. Please do not connect other signal source and pay attention to the

color and label on the ECG leadwires.

- Use only parts and accessories supplied with the patient monitor. Use of any

other accessories may result in out-of-specification performance and possible

safety hazards.

- The monitor must be repaired or disassembled by a qualified biomedical

engineering person.

- DEFIBRILLATOR PRECAUTIONS – Patient signal inputs labeled with the

CF and BF symbols with paddles are protected against damage resulting from

defibrillation voltages. To ensure proper defibrillator protection, use only the

recommended cables and leadwires.

- The monitor is intended for use by persons trained in professional health care.

The operator must be thoroughly familiar with information in this manual

before using the monitor.

- It is the operator’s responsibility to set alarm limit appropriately for each

individual patient.

- If the accuracy of any measurement is in question, verify the patient’s vital

signs by an alternative method and then check the monitor for proper function.

- Set up the monitor in a location for sufficient ventilation. The ventilation

openings of the device must not be obstructed. The ambient conditions

specified in the manual must be always ensured.

- Keep the operating environment free of dust, corrosive, or flammable

materials, and extremes of temperature and humidity.

- Do not operate the patient monitor if it is damp or wet because of condensation

or spills. Avoid using the equipment immediately after moving it from a cold

environment to a warm, humid location.

- Never use sharp or pointed objects to operate the front-panel keys.

- MAINTENANCE – Regular preventive maintenance should be carried out

by a qualified biomedical engineering person annually ( Technical Inspections).

You are responsible for any requirements specific to your country.

- BATTERY POWER – The device is equipped with a battery pack. The battery

discharges even when the device is not in use. Store the device with a fully

charged battery and check the battery status of devices in storage once a month.

- DISPOSABLES – Disposable device are intended for single use only. They

should not be reused as performance could degrade or contamination could

occur.

- DISPOSAL – At the end of its service life, the monitor, as well as its

accessories, must be disposed of in compliance with local governing

ordinances and recycling instructions.

- The built-in battery must be disposal separately according national laws after

the useful life of the device.

- INSTRUCTIONS FOR USE – For continued safe use of this equipment, it is

necessary that the listed instructions are followed. However, instructions listed

in this manual in no way supersede established medical practices concerning

patient care.

- LOSS OF DATA – Should the monitor at any time temporarily lose patient

data, the potential exists that active monitoring is not being done. Close patient

observation or alternate monitoring devices should be used until monitor

function is restored.

- Do not autoclave, ethylene oxide sterilize, or immerse the monitor in liquid.

Unplug the monitor before cleaning or disinfecting the monitor.

- NEGLIGENCE – The manufacturer does not assume responsibility for damage

to the equipment caused by improperly vented cabinets, improper or faulty

power, or insufficient wall strength to support equipment mounted on such

walls.

- 2.4.3 NOTES

- Put the monitor in a location where you can easily see the screen and access the

operating controls.

- This monitor is protected against the effects of cardiac defibrillator discharges

to ensure proper recovery, as required by test standards.

- Hazards arising from software errors have been minimized. Hazard analysis

was performed to meet EN1441:1997.

- Each input and output connection of the monitor is electrically isolated.

- Performance and safety test data are available upon request.

- The equipment is suitable for connection to public mains as define in

CISPR 11.

- 2.4.4 Reference Literature

- Medical Device Directive 93/42/EEC

- EN 60601-1:1990 + A1:1993+A2:1995:Medical electrical equipment general

requirements for safety

- EN 60601-1-1:1994 + A1:1995: General requirements for safety. Requirements

for the safety of medical electrical systems.

## 2.5 Safety Symbols

NOTE:  Some symbols may not appear on all equipment.

Type BF Applied Part：Insulated(floating)applied part is not suitable for

cardiac application. “Paddles” outside the box indicate the applied part is

defibrillator proof.

Medical Standard Definition: F-type applied part (floating/insulated)

complying with the specified requirements of EN 60601-1 Medical

Standards to provide a higher degree of protection against electric shock than

that provided by Type B applied parts.

Type CF Applied Part：Insulated(floating)applied part is suitable for

direct cardiac application. “Paddles” outside the box indicate the applied part

is defibrillator proof.

Medical Standard Definition: F-type applied part (floating/insulated)

complying with the specified requirements of EN 60601-1 Medical

Standards to provide a higher degree of protection against electric shock than

that provided by Type BF applied parts.

Type of protection against electric shock: Class II equipment.

Attention：Consult accompanying documents.

CE0123  The products comply with the requirements of 93/42/EEC, which is enacted

by European Community.

## 2.6 Classifications

UT4000A Patient Monitor is classified, according to EN 60601-1 as:

· Type of protection against electric shock: II, internally powered equipment.

• Degree of protection against electric shock: BF – NIBP, SpO2;

CF – ECG.

• Degree of protection against harmful ingress of water: Ordinary Equipment.

(enclosed equipment without protection against ingress of water)

• Degree of safety of application in the presence of a flammable anesthetic mixture

with air or with oxygen or nitrous oxide: Not suitable.

• Working mode: Continuous running mode.

*  II: Class II equipment;

BF: Type BF applied part;

CF: Type CF applied part;

Not suitable: Equipment not suitable for use in the presence of a flammable

anesthetic mixture with air or with oxygen or nitrous oxide.

## 2.7 Periodic Safety Checks

The following safety checks should be performed at least every 12 months by a qualified person who has adequate training, knowledge, and practical experience to perform these tests.

·  Inspect the equipment and accessories for mechanical and functional damage.

·  Inspect the safety relevant labels for legibility.

·  Inspect the fuse to verify compliance with rated current and breaking

characteristics.

·  Verify that the device functions properly as described in the instructions for use.

·  Test the protection earth resistance according IEC 601-1/1990: Limit 0.1 ohm.

·  Test the earth leakage current according IEC 601-1/1990: Limit: NC 500 uA, SFC

1000uA.

·  Test the patient leakage current according IEC 601-1/1990: Limit: 100 uA (BF),

10 uA (CF).

·  Test the patient leakage current under single fault condition with mains voltage on

the applied part according IEC 601-1/1990: Limit: 5 mA (BF), 50uA (CF).

The leakage current should never exceed the limit. The data should be recorded in an equipment log. If the device is not functioning properly or fails any of the above tests, the device has to be repaired.