# SECTION 3  Basic Description

## 3.1  Front Panel

The front panel of the UT 4000A multi-parameter patient monitor is as follows:

(1)      (2)

(3)(4)   (5)(6) (7)(8)     (9)

| - Handle - Display (3) Power Switch (4) Power Indicator LED (5) (6) (7) (8) (9) Mouse Knob | This handle is convenient for operator to carry the monitor. 5.6 inch TFT LCD screen for displaying waveform, menu, alarm status and vital sign measurement results. Press this key for 1.5 seconds to turn on the patient monitor. Press again to turn off the monitor. When the patient monitor is powered on, the green LED is on. When the patient monitor is connected to AC power supply but not turned on, the yellow LED is on, indicating that the internal battery is being charged. WAVEFORM FREEZE/RESTORE: When a waveform sweep is displayed, press the key to freeze the waveform. When a waveform is being frozen, press this key to unfreeze the waveform and restore waveform sweep. NIBP MEASURE START/STOP: Press this key to start a blood pressure measurement. If the patient monitor is in the process of blood pressure measurement, press this key to stop blood pressure measurement. SILENCE: Press this key to enable or disable Alarm sound. ECG MENU: Press this key to enter ECG and NIBP setup menu. When menu is displayed, rotate the knob to select the functions. When it is turned into pink, press the knob to select the function. Then rotate the knob to select the desired item. In the end, press the knob again to confirm the selection. |
| --- | --- |

## 3.2  Rear Panel

The rear panel of UT 4000A  is as shown in Figure 3-2.

▍▍▍▍▍▍▍▍▍▍▍▍▍▍

2                     5

3

1

4

### Fig. 3-2  Rear Panel Diagram

(1) Idle                                                                    (2) Idle

- RS422 Communication port                             (4) External power supply socket

(5) Heat Exhaust

RS422 communication port can only be connected with GOLDWAY central station.

##### Power Supply Adaptor – there is a power supply adaptor between the patient monitor and  the power cord. The power  supply adaptor consists of a tabletop box with a

##### permanently attached cable and a separate AC line cord. The attached cable is inserted

##### into its mating connector (4 in Fig.2-2)located on the rear of the monitor. Insert the

##### AC line cord into the three-pronged IEC receptacle on the adaptor. The other end of

##### the AC line cord is then inserted into an appropriate wall outlet.

Accessory equipment connected to the analog and digital interfaces must be certified according to the respective IEC standards ( e.g. IEC 950 for data processing

equipment and IEC 601-1 for medical equipment ). Furthermore, all configurations

shall comply with the valid version of the system standard IEC 601-1-1. Everybody

who connects additional equipment to the signal input part or signal output part

configures a medical system, and is therefore responsible that the system complies

with the requirement of IEC 601-1-1. If in doubt, consult the technical service

department or your local representative.

The definition of RS422 port is as follows,

1  N.C. （No Connected）  4 TXD+ (B) （Signal Transmited+）

2  N.C. （No Connected）  5 TXD- (A) （Signal Transmited -）

3  N.C. （No Connected）  6 RXD- (A) （Signal Received -）

7  N.C. （No Connected）  8 RXD+ (B) （Signal Received +）

## 3.3  Side Panel

There are three sockets/receptacles on the left side panel, respectively:

SpO2        This socket accepts SpO2 extension cable.

NIBP        This receptacle accepts the hose connector plug from the blood pressure cuff.

ECG         This socket accepts ECG cable accompanying the patient monitor.

Please consult the specific items in the parameters’ chapters.

## 3.4 Control Keys and Buttons

This patient monitor has three kinds of keys and buttons:

    - Function Key

ECG MENU

SILENCE

NIBP MEASURE START/STOP

WAVEFORM FREEZE/RESTORE

| 3.4.2 Mouse Knob | Located at the lower right of the front panel, corresponding to a combination of the left/right shift key (rotate the knob) and confirm key (press the knob). |
| --- | --- |

3.4.3   Menu key

When rotate the knob, the menu keys as follows will be displayed in the menu area.

ECG Setup；

Trend Display；

System Setup；

Volume Adjustment；

Screen Brightness Adjustment.

## 3.5  Operation

3.5.1 Rotate the knob, select menu key and press the knob to enter the corresponding

Setup menu.

3.5.2 Operation of Setup: Rotate the knob, when the color of a certain key changes from green to pink, press the knob and confirm, then rotate the knob to select the desired option and confirm the selection by pressing the knob.

3.5.3 In the end, when all the setup is completed, Rotate the knob to BACK, press the knob, you can return to the main screen.

3.5.4 The volume can be adjusted by menu key,monitor tone’s volume is adjusted first, then alarm tone’s volume, in the following order:

Rotate Knob  Press Knob   Rotate Knob  Press Knob 

 Rotate Knob Press Knob

3.5.5 The screen brightness can be adjusted by press knob, in the following order:

Rotate  Knob  Press KnobScreen Brightness turn down Press

KnobScreen Turn Bright

The main operator control is the Trim Knob control. The Trim Knob rotates in either direction to highlight parameter labels and menu options. After highlighting the desired

selection, press the knob to view a new popup menu or select.

Remember, when using the Trim Knob control, rotate to highlight, then press to select.

## 3.6 Installation and Connection

WARNING: Before using the monitor for the first time, please read the information

given in chapter 2 “SAFETY”.

· Choose a location which affords an unobstructed view of the monitor screen and easy

access to the operating controls.

· Set up the device in a location which affords sufficient ventilation. The ventilation

openings of the device must not be obstructed (by external equipment, walls or

blankets, for instance). The ambient conditions specified in the Technical

Specifications chapter must be ensured at all times.

· UT4000A is designed to comply with the requirements of  IEC 60601/EN 60601.

· Using the power cord supplied with the monitor, connect it to the AC power

connector. Use only the original cord or an equivalent one.

· Turn on the device with the power switch at the front. The LED is lighted in green,

indicates that the monitor is connected to the AC power.

·  The monitor display appears after approximately 10 seconds.

WARNING:  UT4000A must  be connected to a properly installed power outlet with protective earth contacts only. If  the installation does not provide for a protective earth conductor, disconnect the monitor from the power line and operate it on battery power.

## 3.7 Display illustrations

This patient monitor displays the monitoring results as shown in Fig.3-3:

Fig. 3-3 Screen Display 1

Press the knob,the screen is turned to be as shown in Fig.3-4:

Fig. 3-4 Screen Display 2

The screen display clear vital sign readings, large alarm status, two-channel waveform and data display (ECG waveform, SpO2 waveform and Heart rate/Pulse rate, SpO2 , NIBP  Readings).

At the bottom line of the ECG waveform channel, a time mark appears every second, forming a time scale, which is convenient for observation, calculation and diagnosis.

| WAVEFORM/DATA AREA |  |  |
| --- | --- | --- |
|  |  | STATUS BOX |
| PARAMETER AREA |  |  |
| MENU |  |  |

    - WAVEFORM/DATA  AREA

In this area, the monitor changes from  WAVEFORM AREA to DATA  AREA only when press the knob.

As shown in Figure 3-3, Waveform channel 1 displays ECG waveform; Waveform channel 2 displays SpO2 waveform.

DATA AREA displays NIBP reading set as shown in Figure 3-4.

NIBP reading set includes Time, HR/PR, SpO2, SYS/DIA.

This system is capable of storing up to 200 sets of blood pressure measurement results, but in this channel only 8 sets of readings can be displayed each time. The status bar at the left side of the reading indicates the current position of the reading. Rotate the mouse knob to browse the pages one by one.

NOTE:  If ECG or SpO2 waveform channel displays “No signal received”, please turn off the monitor and turn it on again. If the message still remains, please contact the dealer.

3.7.2   MENU AREA

When rotate the knob, the menu signals will be displayed in the menu area.

ECG Setup；

Trend Display；

System Setup；

Volume Adjustment；

Screen Brightness Adjustment.

3.7.3  PARAMETER  AREA

In the Parameter area, the following parameters are displayed in real time: Heart Rate/Pulse Rate, SpO2, NIBP(SYS/DIA/MAP).

When “ECG” of HR/PR SELECT is selected in the ECG setup menu bar, the

Parameter area will display heart rate reading first. When ECG is not detected, Pulse Rate will be displayed here. When neither heart rate nor pulse rate is detected, the parameter area displays  “___”.

When “SpO2” of HR/PR SELECT is selected in the ECG setup menu bar, the

parameter area will display SpO2 reading first. When SpO2 probe is not connected, heart rate reading is displayed here. When neither heart rate nor pulse rate is detected, the parameter area displays “___”.

NIBP parameter is displayed in mmHg or kPa which can be selected by operator, while “Adult/Pediatric/Neonate” and  “Manual/Auto/Stat” are indicated behind “NIBP”.

3.7.4  STATUS  BOX

In the status box, alarm status and battery power status are displayed.

- Alarm Status:        is used to indicate the alarm sound enable and disable

status.

- Battery Power Status: The battery power status is indicated by the black

segments in The more the segments, the more the battery power.

When it is blank in the frame, the signal     is turned into red and flashes with alarm sounds, indicating the battery power is to be used up. If the battery is not recharged within 3-10 minutes, the patient monitor will be automatically powered off.

3.7.5  WAVEFORM  SPEED

It is displayed on the first row and indicating the sweep speed of vital sign in mm/s.

User may select 12.5, 25 or 50mm/s.

3.7.6  TIME

It is also displayed on the first row and indicating the current monitoring time:

year-month-day, hour:minute:second.

## 3.8 System Setup

The system setup means the monitor system setup, which includes setting up alarm limits and the current time, selecting work mode and wave scanning mode.

Rotate Knob Press Knob

| Alarm Work Mode Monitor | Upper Limit | Lower Limit |
| --- | --- | --- |
| HR/PR Alarm Mode Standard | 120 | 40 |
| SpO2 | 100 | 90 |
| SYS Wave Mode Standard | 160 | 100 |
| DIA | 100 | 60 |
| SYS-DIA Default Restore | --- | 20 |

02 – 03 – 03     12：00：00

EXIT

Alarm limits setup and alarm mode selection is explained in the section 4.

3.8.1  Time Setup

The monitor displays the current time and date. Every time when the monitor is turned on, the system will display the time and date on the first row. To set up the time:

Rotate Knob   Press Knob  Rotate Knob 

TIME  Press Knob Rotate Knob  Press Knob

Rotate the knob to change the day, month, year, second, minute and hour. The new information will not be updated in the system until the knob is pressed.

3.8.2  Work mode

This demo program is for user training only.

WARNING: Never use this program during monitoring patient.

To enter Demo program:

Rotate Knob   Press Knob  Rotate Knob 

MONITOR/DEMO  Press Knob Rotate Knob Press Knob

When the monitor is in Demo mode, you will see  “DEMO” in the screen. Repeat the above procedure to return to normal monitoring mode.

    - Wave scanning mode

To Select single or double scanning line:

Rotate Knob   Press Knob  Rotate Knob 

SINGLE/ DOUBLE  Press Knob Rotate Knob Press Knob

| In the end, after system setup is completed, rotate the knob to BACK , | press the knob |
| --- | --- |

to confirm and EXIT.

## 3.9 Accessories

Use only accessories supplied with the patient monitor. Use of any other accessories may result in out-of-specification performance and possible safety hazards.

All accessories supplied by the manufacturer are listed as follows,

| ### Name | Manufacturer | ###### Type | P/N |
| --- | --- | --- | --- |
| ##### Cuff | Welch Allyn | Adult, Durable | F0222 |
| ##### Cuff | Welch Allyn | Child, Durable | F0212-0216 |
| ##### Cuff | Welch Allyn | Newborn, Single use | F0217-0221 |
| SpO2 Probe | BCI | 3044 Finger Probe | F0414 |
| SpO2 Probe | BCI | “Y” Probe for Infant | F0412 |
| SpO2 Probe | BCI | Wrap Probe for Neonate | F0413 |
| ECG Cable | Changke | SMD 649A | F0153+0154 |
| ECG Electrode | CONMED | REF 1700-030 | F0101 |

## 3.10  Power

3.10.1 AC  Power

UT4000A will be powered by pressing the power button when using AC power. Before applying power, be sure the power requirements match your power supply.

Refer to the label on the back panel of the monitor for the voltage and current requirements.

When all cables are properly connected, press the power button to turn the monitor on. After approximately 10 seconds, you should see a display on the screen. The LED is lighted in green.

3.10.2 Battery Power

The monitor has a built-in battery to provide power to the monitor whenever AC power

is interrupted.

You must charge the battery before using it. There is no external charger. The battery is

charged when the monitor is connected to AC power. A fully depleted battery will take

8 hours when the monitor is in use, to fully charge. To assure a fully charged battery which is ready for use, we recommend that the monitor be plugged into AC power

whenever it is not in use. The LED is lighted in yellow when the monitor is connected to AC power but not in use.

The monitor can work for 1.5 hours on a new, fully charged battery. NIBP monitoring

and the usage of the recorder will drain battery power faster than other parameters.

NOTE:  The battery signal which is at the bottom of the screen flashes with alarm sounds, indicating the battery power is to be used up. If the battery is not recharged within 5 minutes, the patient monitor will be automatically powered off. You should connect

the monitor to an AC power source as soon as you can.

CAUTION: The battery discharges even when the device is not operating.

     Please ensure that the battery is always fully charged when you are keeping the device

in storage for an extended period of time.

     Check the battery status at least once every month and recharge the battery.

3.10.3 Power shut off

When the supply mains to the monitor is interrupted for more than 30 seconds, the

built-in battery will run automatically. If the battery’s power is used up, the monitor will

save the current working status and responding information, relax the pressure in the cuff to undertake the patient’s safety.