SECTION 4 – ALARM CONTROL

## 4.1 Introduction

Alarm is to alert user when abnormal data occur in a monitoring process. It includes audible and visual alarm.

Audible alarm: When monitoring parameter is outside the preset limits, the monitor will automatically activate audible alarm.

Visual alarm: When a monitoring parameter is outside the preset limits, the parameter will automatically flash in the parameter area.

Operator may pull off  the probe from the monitor to test if the visual and audible alarms work normally.

The alarm lever is devided into three classification:

- Crisis: ASYSTOLE,  SYS-DIA is too low; alarm tone is 5 beep;

- Warning: Patient condition exceeds parameter limits, Abnormal Status alarm; alarm tone is 3 beep;

- Message: Battery is low-power; alarm tone is 1 beep.

NOTE:  When “Asystole” is displayed on the screen, please check for the ECG Gain to see if it is too low to detect for its heart rate, if so, user can switch the ECG lead.

Abnormal Status are as follows,

LEAD OFF;                          PROBE OFF;                          LOOSE CUFF;                     AIR LEAK;                          WRONG POSITION;            RANGE EXCEEDED;        OVERPRESSURE;              FAIL;                                      OTHER ERROR;

SEARCH OUT.

WARNING:  After connecting the monitor to the central station, verify the function of the alarm system.

## 4.2 Alarm Setup

To effectively control the system alarm function, the monitor has alarm ON/OFF, alarm mode selection which can be selected through key or menu setup.

4.2.1 Alarm ON/OFF

Press the silence key on the front panel of the monitor, the system may turn on the alarm sound  and set alarm silence within a certain period of time. The system alarm sound ON/OFF status is displayed in the Status Box of the screen, respectively:

: Alarm sound is enabled.

:  Alarm sound will be silenced for 1 minute. Time is counted down. If the

selected alarm mode is “Standard mode”, any new alarm will sound and the

silence command is cancelled.

: Alarm sound will be silenced for 2 minutes. Time is counted down. If the

selected alarm mode is “Standard mode”, any new alarm will sound and the

silence command is cancelled.

| ##### NOTE: Press the | SILENCE | key on the front panel, and at the same time, |
| --- | --- | --- |

- pay attention to the time display under the alarm sign in the Status Box until it is switched to the desired time duration of silence.

- The alarm sound is valid to all monitoring parameters.

- In the stage of silence, if the selected alarm mode is “Standard mode”, any new alarm will sound and the silence command is cancelled. The alarm tone is the new alarm tone in the first 10 seconds, then transfers recording to the lever of alarm.

- Low power alarm isn’t controlled by the silence key, whenever the battery’s power is low, the monitor will alarm instantly.

- 4.2.2 Alarm Limits

- Alarm limits are the basis for the monitor to judge if a parameter is normal or not. There are upper and lower alarm limits. To be higher than upper limits or lower than lower limits are both abnormal and can activate alarm.

- Doctors can set the upper and lower limits according to patient’s status.

| Parameter | Limit Setup Range |
| --- | --- |
| HR (bpm) SpO2 (%) NIBP SYS NIBP DIA SYS-DIA | 20~ 254 50~100 30 ~ 254 mmHg / 4.0 ~ 33.9 kPa 10 ~ 220 mmHg / 1.3 ~ 29.3 kPa 0 ~ 40 mmHg / 0 ~ 5.33 kPa |

- Rotate  the knob, when points to   press the knob to enter the system setup menu.

- Rotate the knob to select the desired parameter (including alarm upper limits and lower limits), and then press the knob to confirm the selection. Rotate the knob left or right to increase or decrease the value until the desired value is obtained. In the end, press the knob to confirm the value.

- In each set of parameters, the data at upper  position are upper limits of alarm, and the data at the lower position are lower limits of alarm.

- When  the upper and lower limits of alarm are set, the system will retain the setting until the next change.

- Even when the monitor is powered off, the alarm limits will still remain in the memory.

4.2.3 Alarm Volume

The volume can be adjusted by menu key, monitor tone’s volume is adjusted first, then alarm tone’s volume, in the following order:

Rotate Knob          Press Knob   Rotate Knob  Press Knob 

 Rotate Knob Press Knob

4.2.4 Default Alarm Limits

After having respetitively set the parameter alarm limits, user may hope to return to the original alarm settings of factory, that is the default setting. Now the steps are as follows:

Rotate Knob    Press Knob  Rotate Knob Default Restore  Press

Knob

The default settings of the monitor are as follows:

| Parameter | Upper Limit | Lower Limit |
| --- | --- | --- |
| HR (bpm) SpO2 (%) NIBP SYS NIBP DIA | 120 100 160mmHg/21.3kPa 100mmHg/13.3kPa | 40 90 100mmHg/13.3kPa 60mmHg/8.0kPa |

4.2.5 Alarm mode

The monitor has two alarm mode, one is standard mode, one is auto mode.  The setup steps are as follows:

Rotate Knob    Press Knob  Rotate Knob Alarm mode  Press knob

 Rotate Knob Standard/Auto Press Knob

- Standard alarm mode

When the abnormal status happen, if alarm is on, all alarms are enable and keep alarming until someone response. The priority and inermission time coincidence corresponding standard.

- Auto alarm mode

Under the following circumstances of LEAD OFF, PROBE OFF, if alarm is on, the system will automatically alarm for 30 seconds and corresponding messages will be displayed.

When blood pressure value is abnormal, and alarm is on, the monitor will alarm for 30 seconds and the abnormal parameter will continue flashing.

If abnormal HR/PR or SpO2 occurs (exceeding the upper or lower limits), and alarm is on, the monitor will activate real-time audible and visual alarm.

In the end, when all the alarm settings are completed, rotate the knob to EXIT  key, press

the knob to confirm the alarm setting and quit the alarm setup menu.