SECTION 5 – ECG MONITORING

## 5.1 Preparation

The ECG monitor continuously monitors the ECG waveform and the pulse. Absence of

“QRS” detection for more than 4 seconds triggers the ASYSTOLE patient alarm. The

LEADS OFF message is displayed if any of the leads are not attached to the patient or ECG cable.

The ECG cable connector is insulated and it is defibrillator-proof.

Use only the cables recommended by manufacturers. These cables ensure patient

safety and protect the device during defibrillation and electrosurgery.

WARNING: Do not come into contact with patients during defibrillation, and be sure that the patient cable is not in touch the metal parts. All of the leadwires should be full inserted in the cable during defibrillation. Otherwise serious injury or death could result.

5.1.1 Skin Preparation

The quality of ECG information displayed on the monitor is a direct result of the quality of the electrical signal received at the electrode.

Proper skin preparation is necessary for good signal quality at the electrode. A good signal at the electrode provides the monitor with valid information for processing the

ECG data.

Following is a suggested guideline for skin preparation:

- Shave hair from skin at chosen sites.

- Gently rub skin surface at sites to remove dead skin cells.

- Thoroughly clean the site with alcohol or a mild soap and water solution. Be sure to

Remove all oil residue, dead skin cells, and abrasives.

- Dry the skin completely before applying the electrodes.

5.1.2 Monitoring Preparation

- Attach ECG electrodes to the cleaned skin, the electrodes’ correct placement is as

shown in Fig. 5-1 ( RA, LA, RL, LL, V).

Fig. 5-1 ECG Electrodes’ Placement

- Connect the ECG leads to the electrodes, then connect the leads to the ECG cable.

Ensure the leads are in the correct position. The ECG leads and patient cable connector

are color-coded according to the AAMI standard for ECG leads. Refer to the ECG

5-lead Electrode Identification table as follows,

| ELECTRODE LABEL | LOCATION | COLOR |
| --- | --- | --- |
| RA | Right Arm | White |
| RL | Right Leg | Green |
| LA | Left Arm | Black |
| LL | Left Leg | Red |
| V | 4th Intercostal Space | Brown |

- Verify the V-lead label is correct if using a 5-leadwire patient cable.

- Adjust ECG setup, if necessary. Follow detailed procedures within this chapter.

5. Setup the ECG alarm limits, please refer to Section 4 Alarm Control.

NOTE：

- The signal input is a high-insulation port and it is defibrillator-proof. The insulated

input ensures patient safety and protects the device during defibrillation and

electrosurgery.

- The ECG patient cable uses a standard AAMI connector. Use only standard AAMI

ECG wires. Do not connect any other signal source to the patient cable.

- Use only high quality silver-chloride electrodes or equivalent which have a negligible

offset potential (offset potential due to polarization should not exceed 130mV for optimal performance).

- Leakage current is limited internally by this monitor to less than 10 uA. However, always consider additional leakage current that can be used by other equipment used

on the patient at the same time as this monitor. So the leakage current test must be done by a qualify engineer before connection other equipment.

- Avoid electrosurgery burns at monitoring sites by ensuring proper connection of the

electrosurgery return circuit. If improperly connected, some electrosurgery units may

allow energy to return through the electrodes.

- Conductive parts of electrodes and associated connectors for applied parts, including

the NEUTRAL ELECTRODE, should not contact other conductive parts including

earth.

- Rate meters may continue to count the pacemaker rate during occurrence of cardiac

arrest or some arrhythmias. Do not rely on rate meter alarms. Keep pacemaker patients

under close surveillance.

- It is normal for the ECG baseline to shift slightly when the monitor is used on cardiac

paced patients.  The monitor displays pacemaker signals exactly as they are

captured. There is no option to display symbolic pace indications.

WARNING:

DO NOT use unprotected lead wires having exposed conductors at the cable end. Use of unprotected lead wires and cables may pose an unreasonable risk of adverse health consequences or death.

Disposable products are intended for single-use only. DO NOT attempt to re-use these products.

## 5.2 ECG  Setup

Rotate Knob Press Knob

Or  press ECG MENU directly, the screen display is as follows:

| ECG LEAD |  |  | Lead selections include I / II/ III/ V/ aVR/ aVL/ aVF/ GND. Default is II. |
| --- | --- | --- | --- |
| WAVE SPEED |  |  | ECG waveform speed selections include 12.5/ 25.0/ 50.0mm/s. Default is 25.0mm/s. |
| ECG SIZE |  |  | Selections include X1/4, X1/2, X1, X2, X4, Auto. Default is X1. |
| The ECG lead and scale selected will be displayed both on the top of the ECG waveform channel. |  |  |  |
| Prior Disp. |  |  | When “ECG” of Prior Disp. is selected, the Parameter area will display HR(Heart Rate) reading first. When ECG is not detected, Pulse Rate will be displayed here. When neither heart rate nor pulse rate is detected, the parameter area displays “_ _ _”. When “SpO2” of Prior Disp.is selected, the parameter area will display PR(Pulse Rate) reading first. When SpO2 probe is not connected, heart rate reading is displayed here. When neither heart rate nor pulse rate is detected, the parameter area displays “_ _ _”. |
| 1mV CAL |  |  | When triggered during ECG monitoring, a 1mV calibration square wave will overlap on the ECG waveform. |

To manually select ECG LEAD, follow this procedure:

ECG MENU Rotate Knob ECG LEAD  Press Knob  Rotate Knob

I / II /III / aVR/ aVL/ aVF/ V/ GND  Press Knob

The monitor  obtains the respiration signal through the same electrodes used for ECG. The ECG lead selection does not affect the respiration signal. The ECG electrodes’ correct placement method is as shown in Fig. 5-1.

Setup the alarm limits of  respiration rate, please refer to Section 4 Alarm Control.

Operation of ECG Setup: Rotate the knob, when arrow is pointing to a certain key, press the knob, then rotate the knob to select the desired option and confirm the selection by pressing the knob. For example,

ECG MENU Rotate Knob  Prior Disp. Press Knob  Rotate Knob

ECG /SpO2/IBP1/IBP2 Press Knob

## 5.3 Precautions

5.3.1    How to assure ECG signals’ quality?

Regardless of patient age, electrodes should be replaced at least every 48 hours to maintain quality signals during long-term monitoring. Over the course of 48 hours,

The electrodes gel will start to dry out and the adhesive will age. After a long period

of  time, the patient’s sensitive skin also may be irritated by the gel or adhesive

causing discomfort.

    - How to decrease ESU interference?

Thorough skin preparation is very important to help keep ESU interference to a minimum. Please refer to 5.1.1 Skin Preparation.

    - How to prevent skin burns?

To prevent skin burns, apply electrocautery electrodes as far as possible from all other electrodes, a distance of at least 15 cm/6 inch  is recommended.

## 5.4 Warnings and Cautions

5.4.1 WARNINGS

- A pacemaker pulse can be counted as a QRS during asystole in either pace mode.

Keep pacemaker patients under close observation.

- Monitor may continue to count the pacemaker rate during cardiac arrest and some

arrhythmias. Therefore, do not rely entirely on the monitor alarms.

- Do not use unprotected leadwires having exposed conductors at the cable end. The use

of unprotected leadwires and cables may pose an unreasonable risk of adverse health

consequences or death.

- ALARMS – Do not rely exclusively on the audible alarm system for patient

monitoring. Ajustment of alarm volume to a low level or off during patient monitoring

may result in a hazard to the patient. Remember that the most reliable method of

patient monitoring combines close personal surveillance with correct operation of

monitoring equipment.

- INTERFACING OTHER EQUIPMENT – Devices may only be interconnected

with each other or to parts of the system when it has been determined by qualified

biomedical engineering personnel that there is no danger to the patient, the operator, or

the environment as a result. In those instances where there is any element of doubt

concerning the safety of connected devices, the user must contact the manufacturers

concerned for proper operation should be verified with the applicable manufacturer’s

instructions for use, and system standards IEC 60601-1-1/EN 60601-1-1 must be

complied with.

- When using High Frequency surgical equipment, you should put the patient cables and

transducers far from the surgical operation in order to reduce the hazards of burns in

the event of a defect in the HF surgical equipment neutral electrode connection.

- Using the electrodes made of deferent material at the mean time may cause electrode

excess polarization and the monitor can’t work normally. Use only ECG electrodes suggested by the manufacturer.

5.4.2 CAUTIONS

- ECG patient cable shall be connected to the ECG leadwire supplied with this monitor. Please do not connect other signal source and pay attention to the color and label on the ECG leadwires.

- Use only parts and accessories supplied with the patient monitor. Use of any other accessories may result in out-of-specification performance and possible safety hazards.

- DEFIBRILLATOR PRECAUTIONS – Patient signal inputs labeled with the

CF and BF symbols with paddles are protected against damage resulting from

defibrillation voltages. To ensure proper defibrillator protection, use only the

recommended cables and leadwires.

4. The ‘LEAD OFF ’ message could be the result of overload (saturation) condition of the ECG input (due to a discharge of a defibrillator). First check ECG electrodes contact with patient’s skin and check all leadwire connections OK.  If possible please use new electrodes and then please wait less than 10 second. If still show ‘LEAD OFF ’ message, please contact with supplier.

## 5.5 Troubleshooting

5.5.1 Inaccurate heart rate and false asystole

- Check ECG signal from patient:

- Check/adjust lead placement.

- Check/perform skin preparation.

- Check/replace electrodes.

2. Check/adjust amplitude of  ECG waveform.

5.5.2 No ECG waveform

When leadwires are connected, there is no ECG waveform. The screen displays “LEAD OFF” or “NO SIGNAL”.

1. Check ECG electrodes contact with patient’s skin and check all leadwire connections.

- No Resistance should exist between extension wires applied to patient body to the

leads of the ECG connector. If it is indefinitely large, it shows that the leadwire is broken and should be replaced.

- If ECG waveform channel shows “No signal ”, then it shows there is some problem in the communication between ECG module and the main system. If this message still remains after the monitor is turned off and on again, please contact the supplier.

5.5.3 ECG Baseline Shift

ECG baseline is not stable, occasionally shifts out of the display area.

- Check if the working environment is very humid and if the monitor has moisture inside. If so, turn on the monitor and let it run for 24 hours to remove moisture inside.

2．Check if the electrode’s quality is good, and the skin surface where the electrode is

applied is clean. If so, replace it with good electrode, and clean the skin surface where

the electrode is applied.