SECTION 6 – NIBP MONITORING

## 6.1 General Information

The monitor’s automatic electronic sphygmomanometer measures and displays a patient’s

arterial blood pressure, mean arterial pressure, and pulse rate using the oscillometric technique. Oscillations are caused by blood pressure pulses against the cuff. When the pressure in the cuff decreases, a sensor located in the monitor detects pressure fluctuations in the cuff. These pressure fluctuations are due to arterial volume changes that result from the blood flow as cuff pressure falls.

The first number of a typical blood pressure reading (systolic) represents the maximum

pressure generated when the left ventricle of the heart contracts. When the ventricles

relax, pressure in the arteries decrease as blood flows out of the arterial system into the

capillary system. The lowest point that the pressure reaches before the next ventricular

contraction represents the second number of the blood pressure measurement (diastolic).

The mean arterial pressure (MAP), which is calculated by the equation 1/3 (systolic+2×diastolic), corresponds to the maximum pulse amplitude at the lowest pressure level.

Note: Because the methods of  Auscultation and  Oscillation are different, one cannot be

used to check the accuracy of the other.

## 6.2 Monitoring Preparation

- Select a cuff appropriate for limb size (refer to 6.3 Cuff Selection & Application).

- Place the cuff properly on patient(refer to 6.3 Cuff Selection & Application).

- Connect patient cable to monitor.

- Avoid compression or restriction of pressure tubes, and make sure the tube between the cuff and the monitor are not kinked or blocked.

- Select correct measure patient (adult/child/neonate) from the NIBP setup menu.

- Confirm that the cuff is fully deflated before positioning on the patient.

- The measurement should be performed on the left arm if feasible. Roll up sleeve

before measurement. Only very thin fabrics will not impair the measurement.

- Check that rolled up sleeve does not constrict the vessels.

- The arm should be placed on a surface that is level with the heart. The palm of the

hand should face up.

10.Position patient so that no external pressure is applied against cuff while monitoring.

External pressure may cause inaccurate blood pressure values.

11.Start  NIBP measurement.

WARNING:

- The measure patient selected in the NIBP setup menu and the cuff size used must

be correct to obtain reliable NIBP data and to prevent overpressure in neonatal

or pediatric use.

- The monitor is not recommended for patients experiencing seizures or tremors, or

patients with hypotension, hypertension, arrhythmias, or extremely high or low

heart rate. The software algorithm cannot accurately compute NIBP on patients

with these conditions.

- Do not place the cuff on a limb being used for A-V Fistulas, intravenous infusion,

or any area where circulation is compromised or has the potential to be compromised.

- Use care when placing cuff on extremity used to monitor other parameters.

- Use only the accessaries suggested by the manufacturer. Otherwise the safety will

be degraded.

- The monitor is coincidence with EMC requirements, but if the electromagnetic interference is too intense, the NIBP measuring may be influenced.

- When continues mode used long time, the skin under the cuff may occur nerve damnified or blood lost. Periodically check patient limb circulation distal to the cuff. Check frequently when using Auto NIBP in 1 and 2 minute intervals and continues mode. Intervals below 10 minutes are not recommended for extended periods of time.

- Operator should take off the cuff from the patient when the power is shut off.

- After monitoring, cuff should be cleaned and sterilized for next use.

10.The equipment is suitable for use in the presence of electrosurgery.

11.When using High Frequency surgical equipment, you should put the patient

tube and cuff far from the surgical operation in order to reduce the hazards of

burns in the event of a defect in the HF surgical equipment neutral  electrode

connection.

The NIBP cable connector is insulated and it is    defibrillator-proof . Use only the cuffs recommended by manufacturers. These cuffs ensure patient safety and

protect the device during defibrillation and electrosurgery.

WARNING: Do not come into contact with patients during defibrillation, and be sure that the patient cable is not in touch the metal parts. Otherwise serious injury or death could result.

## 6.3 Cuff Selection & Placement

The accessories belonging to the pressure measuring system: Durable adult cuff, Durable

child cuff,  Durable Newborn cuff.

It is important to select the proper size blood pressure cuff and place the cuff correctly.

Cuff selection:

- Identify patient’s limb circumference.

- Select appropriate cuff (limb circumference is identified on each cuff ).

Cuff placement:

- Place cuff snugly around extremity being used.

- Marking on cuff should match artery location.

- Cuff should be 2 to 5 cm (1 to 2 inches) above the elbow is using brachial artery.

- The air pad should be exactly over the brachial artery. Tubing is immediately to the

right or left of the brachial artery to prevent kinking when elbow is bent.

## 6.4 NIBP Setup

- NIBP(Non-Invasive Blood Pressure) setup includes: change NIBP units, select Patient (Adult/Pediatric/Neonate) ,change NIBP measurement mode (Auto/Manual/Stat), setup the interval of automatic NIBP measurement.

Rotate Knob Press Knob

Or  Press ECG MENU directly, the screen display is as follows：

- 6.4.1  Patient Selection

- Patient selections: ADULT (Adult)/ PEDI(Pediatric) / NEO(Neonate).

- When performing NIBP measurements on neonates and infants, be sure that the PATIENT Selection is set to be “NEO”.

- For the child under 12 years old, it is recommended to select “PEDI”.

- 6.4.2  NIBP Mode

- NIBP measurement mode:

- MANU(Manual)

- AUTO (Automatic)

- STAT (Continuous)

- The current NIBP mode is displayed at the upper right of the NIBP parameter.

- Under Manual (MANU) NIBP measurement mode, press

| - NIBP MEASURE START/STOP | - key on the front panel, blood pressure measurement |
| --- | --- |

- will start immediately. When blood pressure measurement starts, the monitor first inflates  the cuff to a certain pressure, then slowly releases the cuff pressure and measures the patient’s blood pressure value.

- Under Auto blood pressure measurement mode, the monitor will start inflating at the start of every auto cycle and measuring patient’s blood pressure value automatically. Under Auto mode, if patient’s blood pressure is too high, the monitor will adjust the maximum inflation pressure according to the actual blood pressure level.

- Under Stat blood pressure measurement mode, patient’s blood pressure is measured continuously within five minutes. This mode is used to closely observe patient’s blood pressure changes under a very special situation.

- When blood pressure is measured manually or through pressing keys, the initial pressure value of  cuff is 175mmHg for adults and 95mmHg for neonates. For Auto and Stat modes, the initial inflation pressure is the same as above, and then, based on the previous measurement results, adjust the initial inflation value of auto measurement.

- To change the blood pressure measurement mode:

| - ECG MENU | -  Rotate Knob  | - MODE | -  Press Knob Rotate Knob |
| --- | --- | --- | --- |

| -  | - AUTO/MANU/STAT/TEST | -  Press Knob |
| --- | --- | --- |

    - Measurement Time

- The current blood pressure measurement time is displayed after the measurement mode.

- Total continuous measurement time is displayed after “STAT”: 0:05(5 minutes).

- Automatic measurement time interval is displayed after “Auto”. Selections can be made from 1-10´, 15´, 30´, 60´, 90´, 120´ (´ is for minute, the time is the interval between the start of last blood pressure measurement and the start of next measurement), which can be adjusted by rotating the knob left or right. To adjust the interval:

| - ECG MENU |  | -  Rotate Knob  |  | - Measure Time | -  Press Knob  Rotate Knob |  |
| --- | --- | --- | --- | --- | --- | --- |
| -  | - 0:01/0:02…/2:00 |  | -  Press Knob |  |  |  |

- 6.4.4  Changing NIBP Scale

- The monitor has two blood pressure measurement scale, i.e. mmHg and kPa. Default is mmHg.User may set the blood pressure unit according to the needs.

- The change of the blood pressure scale will be displayed with all the blood pressure parameters, e.g. parameter area and trend graph.

- To change the blood pressure scale:

| - ECG MENU | -  Rotate Knob  | - Pressure Unit | -  Press Knob Rotate Knob |
| --- | --- | --- | --- |

| -  | - mmHg/kPa | -  Press Knob |
| --- | --- | --- |

## 6.5 Troubleshooting

6.5.1 Status Messages

A  status message will be displayed in the NIBP values windows if a measurement is unable to be completed. Following is a list of the status messages with monitor response and action to take.

| ### Status Message | ### Monitor Response | ### Solution |
| --- | --- | --- |
| LOOSE CUFF | Sytem status alarm. | Check cuff’s size and placement. Select proper cuff and place it correctly. |
| AIR LEAK | Sytem status alarm. | Check for leak in cuff or tubing, if one is air-leak, change it. |
| WRONG POSITION | Sytem status alarm. | Check patient and cuff placement, if wrong, correct it. |
| OVER PRESSURE | Sytem status alarm. | Remove cuff and contact service. |
| RANGE EXCEEDED | Sytem status alarm. | Contact service. |
| OTHER ERROR | Sytem status alarm. | Possible excessive patient movement or arrhythmia condition. Check patient or contact service. |

6.5.2 Prevent Erroneous NIBP measurement

To prevent erroneous NIBP measurement:

- Check for proper cuff size, too small a cuff can give an erroneously high value, too

large a cuff can give an erroneously low value.

- Check for residual air left in the cuff from a previous measurement.

- Make sure cuff is not too tight or too loose.

- Make sure cuff and heart are at same level, otherwise hydrostatic pressure will offset the NIBP value.

- Minimize patient movement during measurement.

- Check for leak in cuff or tubing.

- Calibration should be at least taken once a year.

## 6.6  Maintenance

Cuff should be cleaned with a warm and mild detergent solution after each using.

Calibration should be taken by qualified professional personnel once a year.

| - ECG MENU | -  Rotate Knob  | - Calibration | -  Press Knob Rotate Knob |
| --- | --- | --- | --- |

| -  | - -6/-5/-4/-3/-2/-1/1/2/3/4/5/6/Auto | -  Press Knob |
| --- | --- | --- |

The method for calibrating is introduced in technical manual. Qualified professional personnel may refer to the technical manual.