SECTION 7 SpO2 MONITORING

## 7.1 Introduction

SpO2 monitoring is a noninvasive technique used to measure the amount of oxygenated

hemoglobin and pulse rate by measuring the absorption of selected wavelengths of light.

The light generated in the probe passes through the tissue and is converted into an electrical signal by the photo detector in the probe. The monitor processes the electrical

signal and displays on the screen a waveform and digital values for SpO2 and pulse rate.

Setup the alarm limits of  SpO2, please refer to Section 4 Alarm control.

## 7.2 Probes

To ensure conformance with all safety and performance specifications, use only the recommended accessories when performing Oxygen Saturation monitoring. These are available from GOLDWAY.

| - Probes | - Recommended Application |
| --- | --- |
| - Finger Probe for Adults, Reusable - “Y” Probe for Infant, Reusable - Probe for Neonate, Reusable | - Attach the finger probe to the patient .Be sure to - fully insert the patient’s finger into the probe. - Run the cable along the hand. - Attach the “Y” probe to a fleshy portion of the - infant’s hand/foot. Attach the probe with the light source side on the top or outside of the - hand/foot to keep the detector side away from - ambient light. Line up the light source side with the detector side, so the source and the detector are in the same plane. Secure the probe with medical tape, being careful not to over tighten the tape. - Attach the neonatal probe to the fleshy part of the neonate’s hand/foot. Line up the light source side with the detector side, so the source and the detector are in the same plane. Secure the probe and cable with medical tape, being careful not to over tighten the tape. |

The signal input is insulated and it is     defibrillator-proof. The insulated

input ensures patient safety and protects the device during defibrillation and

electrosurgery.

CAUTION: Use only accessories supplied by the manufacturer. Use of any other accessories may result in out-of specification performances.

WARNING: Do not come into contact with patients during defibrillation, and be sure that the patient cable is not in touch the metal parts. Otherwise serious injury or death could result.

WARNING:  Prolonged use or the patient’s condition may require changing the SpO2 probe site periodically. The recommended application time for each probe at a single site is 2 hours. Change probe site and check skin integrity, circulatory status, and

correct alignment at least every 4 years.

WARNING:  In the monitoring of patients the coincidence of adverse conditions may lead to a disturbed signal going unnoticed. In this situation artifacts are capable of

simulating a plausible parameter reading, so that the monitor fails to sound an alarm.

In order to ensure reliable patient monitoring, the proper application of the probe and the

signal quality must be checked at regular intervals.

## 7.3 Precautions

7.3.1 Measuring neonates and infants

The display of inaccurate pulse oximetry (SpO2) values has been linked to the presence of poor signal strength or artifact due to patient motion during signal analysis. This condition is most likely to be encountered when the monitor is used on neonate or infants.

These same conditions in adults do not impact the pulse oximetry values to the same extent.

When measuring neonates and infants, always observe the following precautions:

- The pulse rate determined by the pulse oximetry function must be within 10% of the heart rate, and

- Six seconds of artifact-free waveform must be present, or

- Stable pulse oximetry values are displayed for six seconds.

7.3.2 Probes Application

- Clean the surface of the probe before and after each patient use.

- Following the instructions of the probes, correctly position and attach the probe to your patient.

- Be sure to read all literature accompanying probes for specific safety information.

WARNINGS:

- PATIENT SAFETY - If a probe is damaged in any way, discontinue use immediately.

- DATA VALIDITY – Do not allow tape to block the probe light detector.

- DATA VALIDITY – Do not expose probe detector to strong ambient light while monitoring a patient. A poor signal may result.

7.3.3 Other factors causing inaccurate readings

- Nail polish, especially violets and blues, can also reduce the reading due to the same

absorption. Removal of  the polish is always recommended.

- Long nails and artificial nails can interfere with good  detector opposition through the

the tissue.

- Patients with severe anemia can have low SpO2 readings.

- Significant levels of dysfunctional hemoglobins, such as carboxyhemoglobin or methemoglobin, will affect the accuracy of the SpO2 measurement.

- SpO2 measurements may be adversely affected in the presence of high ambient light. If

necessary, shield the sensor area (with a surgical towel, for example).

- The SpO2 function is calibrated to read functional arterial oxygen saturation.

Significant levels of dysfunctional hemoglobins such as carboxyhemoglobin or

methemoglobin may affect the accuracy of the SpO2  measurement.

NOTE: Do not attach the SpO2 probe to the same limb with a blood pressure cuff.

NOTE: The monitor is calibrated to display functional saturation.

NOTE: The possibility of hazards arising from errors in the software programme is

minimized by improving the reliability of the software.

## 7.4 Troubleshooting

7.4.1 No SpO2 reading

In the process of monitoring, there is no SpO2 waveform and data.

Check the finger probe to see if there is red light, if the arm is being pressed, and if the room temperature is too low.

If there is no red light in the finger probe, check the extension cable and connecting part. When room temperature is very low, do not expose patient’s arm outside , which will affect the measurement results.  Do not measure blood pressure and SpO2 on the same arm. Pressing the arm will affect SpO2 measurement.

If SpO2 waveform channel shows “NO SIGNAL”, it indicates that there is a problem in the communication between SpO2 module and the main system. If the message still remains after turning the monitor off and on again, please change the SpO2 board.

7.4.2 Intermittent SpO2 Signal

During SpO2 measurement, the reading is off and on. Inspection method:

- During prolonged monitoring and surgery, check if patient arm movement results in interruption of SpO2 measurement.

- Check the SpO2 extension cable.

Solution:

Keep the patient steady. It is normal if  SpO2 value is lost due to hand movement. If SpO2 extension cable is faulty, replace it.