# APPENDIX I - TECHNICAL SPECIFICATIONS

## A Safety Approval & Quality System

Designed to meet EN 60601-1:1990 + A1:1993 + A2:1995

Class II Equipment, internally powered equipment, double insulated

Type BF,CF applied parts

ISO9001 & EN46001 Certified

## B Power Requirements

| Power Supply | 110-230 VAC, 50/60 Hz, via power adapter |
| --- | --- |
| Input Power |  80VA |
| Battery Charge time Operating time | DC 12V/2.2AH sealed lead-acid 10 hours 1.5 hours (full recharge) |
| Battery Charging Method | Automatic charging after monitor is connected to AC power supply (with charge protection function) |
| Discharge Protection | When powered by battery, the monitor will be automatically turned off when battery power is almost used up. |

## C Performance Specifications

1 ECG

| Patient Safety Standard | EN 60601-1: 1990 |
| --- | --- |
| CMRR | 60dB (Common Mode Rejection Ratio) |
| Heart Rate Range | 20 ~ 254bpm  1bpm |
| Heart Rate Averaging | 8 second average |
| Interface | AAMI 6-pin |
| Lead Selection | I, II, III, aVR, aVL, aVF, V (5 lead mode) |
| Lead Fault Alarm | Audible, Visual |
| Input | 5-lead ECG patient cable |
| QRS Indicator | Audible and Visual Alert |
| Sweep Speed | 12.5/25/50 mm/sec |
| Gain Selection | x4, x2, x1, x1/2, x1/4,Auto |
| Trends | 12 hours |
| Patient Isolation Frequency range | Breakdown voltage: 4000VAC 50Hz 60 seconds Leakage current: <10A 0.5 ~ 35Hz (+0.4dB,-3.0dB) |
| Patient Drive Current | <10A |
| Enclosure Leakage Current | <0.1mA |
| Maximum T wave Rejection Capability | 1.2mV |
| Heart Rate Alarm Response Time | < 12 seconds |
| Aspect Ratio | 0.24 ~ 0.6 sec/mV |
| Defibrillator Protected & ESIS Protected | Tested with 5kV |
| Recovery Time Following Defibrillation | <5 seconds |

2  Pulse Oximetry ( SpO2 )

Measure Range                       50 ~ 100%

SpO2 Averaging		8 second average

SpO2 Accuracy		2% (70 ~ 100%),  3% (50 ~ 69%)

Pulse Rate Range	            30 ~ 250bpm

Pulse Rate Averaging	            8 beat average

Pulse Rate Accuracy	            1% ( 30 ~ 100bpm)

Sensor Types			Finger, Universal “Y”, wrap probes

Pulse Rate Display		Digital

Refurbish time                        about 2 seconds

Wavelength           Infrared ：940nm

Red ： 660nm

Energy                  Infrared ≤ 22.5 mW

Red ≤ 30 mWv

3  Non-Invasive Blood Pressure ( NIBP )

Method			Automatic oscillometric

Parameters                   	Systolic, diastolic, mean arterial pressure, pulse

Scale				mmHg or kPa

Operating Modes		Manual, Automatic, Continuous

Repeat Cycles 			1 ~ 10, 15, 30, 60, 90, 120 minutes

Measure Range

Systolic          Adult/pediatric	30 ~ 254mmHg (4.0 ~ 33.9kPa)

Neonate    		30 ~ 135mmHg (4.0 ~ 18.0kPa)

Diastolic	Adult/pediatric	10 ~ 220mmHg (1.3 ~ 29.3kPa)

Neonate 		10 ~ 110mmHg (1.3 ~ 14.7kPa)

Cuff  Pressure Range		Adult/pediatric	0 ~ 280mmHg (0 ~ 37.3kPa)

Neonate		0 ~ 140mmHg (0 ~ 18.7kPa)

Initial Cuff  Inflation		Adult/pediatric	17010mmHg (22.71.3kPa)

Neonate		10010mmHg (16.01.3kPa)

Deflation Pressure		30mmHg(4.0kPa) higher than the last systolic pressure.

Cuff Inflation Rate		No greater than 50mmHg/sec

Measurement Time		Typical 25 seconds

Maximum 40 seconds

Typical Stat 20 seconds

Pressure Display Accuracy	3mmHg

BP Pulse Rate Accuracy	2% @ 40 ~ 240bpm.

Cuff				Neonate, infant, pediatric, standard adult.

## D TFT Color Display

Size				5.6″

Matrix				320(H) x 240 (V) pixels

Viewing Angle                       160°

## F Environmental Specifications

Temperature		            Operating	                     +10 ~ +40C

Transport and Storage	         -20 ~ +60C

Relative Humidity	            Operating	                     ≤85% (non-condensing)

Transport and Storage	         10 ~ 90% (non-condensing)

Atmospheric pressure             Operating                              700 ~ 1060 hPa

Transport and Storage          500 ~ 1060 hPa

Specifications subject to change without notice.