# APPENDIX II - MAINTENANCE

## A  Monitor

When necessary, clean the exterior surfaces of the monitor with a soft cloth dampened with a warm and mild detergent solution. Do not allow liquids to enter the interior of the instrument.

- Electrical shock and flammability hazard - always turn the monitor off and disconnect it from AC main power before cleaning.

- You should  wipe off liquid with a dry cloth immediately when accidental wetting of the equipment, do not allow liquid to enter the interior of the instrument. If liquids have entered a device, take it out of service and have it checked by a service technician before it is used again.

- Do not autoclave or pressure sterilize this monitor. Do not immerse this monitor in any liquid. Do not gas sterilize this monitor.

- Do not touch, or rub the display screen with abrasive cleaning compounds, instruments, brushes, rough surfaced materials or make contact with anything that can scratch the panel.

## B  NIBP Cuff and SpO2 Sensor

Clean the cuff, tubes and sensor after each patient use.

If cuff disinfection is required, wipe the cuff’s surfaces with Isopropyl alcohol and use water rinse.

The cuff require sterilization periodically, and the operator should be sure to follow the hospital procedures.

WARNING: Do not sterilize SpO2 sensor by irradiation, steam, or ethylene oxide. Do not immerse in liquids.

CAUTION: Before cleaning or disinfecting the sensor, unplug it from the monitor.

Clean the sensor with a soft cloth moistened with water or a mild soap solution. To disinfect the sensor, rub with isopropyl alcohol. Allow the sensor to dry completely after cleaning.

If a sensor has any sign of wear or splitting, please replace it immediately.

## C  Patient Cable (ECG Cable, SpO2 Extension Cable)

- Do not autoclave the patient cables.

- Never submerge the patient cables in any liquid or allow liquids to enter the electrical connections. Wipe the exterior surface of the cables with soap and water or alcohol.

CAUTION: The decision to sterilize must be made per your institution’s requirements with an awareness of the effect on the integrity of the transducer and dome.

## D Periodic Safety Checks

The following safety checks should be performed at least every 12 months by a qualified person who has adequate training, knowledge, and practical experience to perform these tests.

·  Inspect the equipment and accessories for mechanical and functional damage.

·  Inspect the safety relevant labels for legibility.

·  Inspect the fuse to verify compliance with rated current and breaking characteristics.

·  Verify that the device functions properly as described in the instructions for use.

·  Test the protection earth resistance according IEC 601-1/1990: Limit 0.1 ohm.

·  Test the earth leakage current according IEC 601-1/1990: Limit: NC 500 uA, SFC

1000uA.

·  Test the patient leakage current according IEC 601-1/1990: Limit: 100 uA (BF),

10 uA (CF).

·  Test the patient leakage current under single fault condition with mains voltage on

the applied part according IEC 601-1/1990: Limit: 5 mA (BF), 50uA (CF).

The leakage current should never exceed the limit. The data should be recorded in an equipment log. If the device is not functioning properly or fails any of the above tests, the device has to be repaired.