APGAR V1.0 Circuit Operation.
Note:
Design calculations are detailed in the “Hardware Design Detail” document.

Power supply.
The power requirements were:

To operate from “AA” size batteries,

To operate for a minimum of 12 months before a battery change is necessary,

To achieve a very low power dissipation design.

A high efficiency step-down switch-mode design was chosen to fulfil these requirements.  

The controller used was an LTC1877 step-down converter in a standard design.

An input voltage range of 3.4V – 6.4V, from 4 off ”AA” batteries, is converted to an output of 3.35V.  This output voltage was chosen because of the tolerances in the ST7FLITE15 processor’s Low Voltage Detect (LVD) Low Threshold specification – a maximum of 3.15V.  The LVD circuit is used to reset the ST7 processor at power-up.  So, to achieve minimum power dissipation, 3.35V was chosen for the VDD supply.  A “PP3” battery solution was considered but dropped due to the insufficient capacity of the PP3 battery.

High capacity ceramic capacitors were used for C2 and C4 in preference to tantalum or other electrolytic types, as their technical performance is better.

Processor and peripheral circuitry.
The processor is an ST7FLITE15, a twenty-pin device.

Power supply input.
The VDD and VSS pins are decoupled by a 0.1uF ceramic capacitor to provide noise and ESD pulse filtering.  Bulk capacitance is provided by C2 in the power supply, which is mounted very close to the processor’s power supply input pins on the PCB.

Oscillator.
The oscillator is a 32.768kHz crystal type chosen for accuracy and stability.  C7 and C8 form the load for the crystal.  The processor’s internal 700kHz (at VDD=3V) RC oscillator is not used in this design because of its poor stability and relatively high power consumption.  Because of its higher frequency, it would also radiate more RF interference than the 32.768kHz oscillator will do.

A/D functionality and accuracy check.

The circuit IC3 and R5 is used to check the functionality and accuracy of the analogue-to-digital (A/D) converter.  IC3 is a 1.22V voltage reference.  It is fed from processor pin PB5 via R5.  The reference voltage generated is monitored by processor pin PB4 – the AIN4 A/D input.  When the circuit is not in use, all three processor pins (PB4, PB5 and PB6) are at VSS level to minimise power dissipation.  In use, PB6 is kept at VSS whilst PB5 is raised to approximately VDD.  The reference voltage at PB4 is converted to digital form and compared to a target window of digital values.  If the result is within this window then the A/D converter is considered functional and accurate, otherwise an error is reported.  This test is not foolproof.  It does not use the A/D input which actually monitors the battery voltage, so that input could be faulty and not be detected.

Battery monitoring.
A battery monitoring circuit comprising R3, R4, C9 and D3 feeds the AIN0 input to the processor and the internal analogue-to-digital converter.  R3 and R4 form a potential divider such that when the batteries are new, the 6.4V battery voltage is brought within the A/D input range – less than VDD.  C9 provides noise filtering.  D3 protects the processor AIN0 input from damage in case R4 goes open circuit by clamping the input to one (Schottky) diode voltage drop above VDD.

The battery voltage is checked once a day and compared to a 4V threshold.  If the voltage is below 4V, then a “Lo BAtt” message is displayed on the LCD.

Sounder.
A piezo-electric sounder is attached to processor pin PA4, the PWM2 output pin.  The pulse width modulator controls the frequency and amplitude of the APGAR and key-press beeps.  Diode D2 protects the processor pin from any voltage spikes produced by the sounder.

Keypad inputs.
Processor pins PA3 (Start), PA2 (Freeze) and PA1 (Reset) are the inputs from the press-to-make membrane keypad switches.  Each switch connects an input to VSS when made.  Each pin has an internal pull-up resistor, so an external resistor is not required.  Resistors R14, R15 and R16 provide some protection against any ESD pulses induced into the membrane’s keypads.  The membrane keypads connect to the PCB via CN1.

Programming and debugging inputs.

A six-way connector, CN2, provides the programming and debugging interface to the design.  The pin functions are:

CN2/1
VSS

CN2/2
VDD

CN2/3
Processor reset.  A pull-up resistor, R17, enhances noise immunity.

CN2/4
ICCCLK.  This pin provides the programming/debug clock to the         processor.  A pull-up resistor, R9, provides enhanced noise immunity.

CN2/5
ICCDATA.  This pin provides programming/debug data.  A pull-up resistor, R8, provides enhanced noise immunity.

CN2/6
OSC1.  This is the CPU’s external clock input, if needed.

I2C communications interface.
This is provided by processor pins PB1 (Serial Clock), PB2 and PB3 (Serial Data).  This interface is used to communicate with the LCD controller chip, IC4.  As the ST7FLITE15 processor does not have a hardware I2C interface, all the communication takes place under software control and execution.  All I2C communications are checked, in software, for errors.

Spare processor pins.
Pins PA0 and PA7 are not used and are tied to VSS and VDD respectively, through resistors, to minimise power dissipation.  The resistors provide some ESD protection.

Hardware watchdog.
Currently the processor’s hardware watchdog is not used.

LCD controller and display.
The LCD controller, IC4, is a PCF8576C.  It is a self-contained controller in that it has its own oscillator and LCD voltage converter circuitry on-board.  It communicates with the controlling processor via the I2C bus, R6 and R7 providing the pull-ups for the bus.  The controller is operated in low power mode to minimise power dissipation.  In this application, this mode does of operation does not affect the I2C communication speed.

Only one back plane output (BP0) is used as the LCD is operated in static mode.  The segment drivers were allocated to the LCD pins in a way that made PCB layout easy, the resulting non-linearity being corrected in software.

When the LCD is not displaying the APGAR time, it displays either a flashing colon or the alternating words “Lo” and “Batt”.  If it is the flashing colon, then the LCD controller, IC4, is quite capable of providing the flashing function.  However, it was decided to make the processor do the flashing by alternately sending “:” and “ “.  This was done so that the processor could not “die” without it being noticed i.e. no flashing colon.  The alternating words “Lo” and “Batt” also show that the processor is functioning.

P. Anderson

28th January 2005

PAGE  
1

