	     [image: image4.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley - West Yorkshire, BD20 7DT - United Kingdom

Tel: +44 (0)1535 634542         Fax: +44 (0)1535 635582

Email: info@viamed.co.uk       www.viamed-online.com



SpO2 Clinical Study Report
Confidential – Distribution only by written permission of Viamed Ltd.

SpO2 Clinical Study Summary Information

Type: VM-2105 (0012105)

Sensor Tested: 
VM-2105 (integrated sensor)

Test Date(s): 
25 – 26 Aug 2009 
Monitor Tested: VM-2105; S/N 30000323 (grey)

Report Date:
28 Aug 2009 

Data Analysis

SpO2 and Co-oximeter data were inputted into Excel for analysis.

Functional SaO2 (percent) was calculated as SaO2 (%) = O2Hb / [100 – (COHb + MetHb)].

Two co-oximeters were used: OSM3 for SaO2 values; ABL520 to check validity by agreement.

Bland-Altman statistics; Bias, Standard Deviation of Difference (SpO2 – SaO2).

Regression Line: Slope, Intercept, and Correlation Coefficient.

RMS Error: Calculated per ISO 9919 as: 
[image: image1.jpg]. \/z (SpO, - S’

RMS —





Test Results

Refer to the graphs on the following page:

	Samples (n)
	220
	Subject Demographics

	Bias:
	-0.177
	Male:
	60%

	Std. Dev.:
	1.513
	Female:
	40%

	RMS Error:
	1.5194
	Age Range:
	20-37

	Slope:
	1.041
	Dark Skin:
	30%

	Intercept:
	-3.725
	Light Skin:
	70%

	Correlation (R2)
	0.967
	Excluded Points:
	None (see page 2)


Protocal Information

Protocol: (title): Validation of Pulse Oximeters
Investigator: Dr. Philip S. Clifford, Ph.D.

Review Committee Approval: VA Project No: 1445-05 HRRC No.: 73-93

Facility: Medical College of Wisconsin, VA Medical Center, Milwaukee, WI 53295 USA

Subjects: 10 subjects per test series, minimum 20% dark skin

Number of samples: ≥ 20 SpO2 – SaO2 pairs/subject with co-oximetry in the range 70% - 100%

Blood Sample: via radial artery

Reference Method: Co-oximeter(s) – Radiometer OSM3 reference with ABL700 Co-oximeter check.
Subject Consent: Prior written informed consent was obtained for all subjects.
Subject Confidentiality: Confidentiality was maintained for all subjects per number system.

Protocol Execution: Performed in accordance with ISO 9919 Annex EE, analysed to (1), (2).

Technical References

(1) ISO 9919:2005 (IEC 60601-2-54): Medical electrical equipment – Particular requirements for the basic safety and essential performance of pulse oximeter equipment for medical use.

(2) (2) BLAND, J.M., ALTMAN, D.G. Statistical methods for assessing agreement between two methods of clinical measurement. Lancet. (8 Feb), (1986), pp. 307 – 310.

Conclusion

The tested device meets the stated accuracy specification for RMS Error ARMS = 1.5
for the range 70% - 100% SaO2. The accuracy is not specified below 70% SaO2.

SpO2 Clinical Study Report

Confidential – Distribution only by written permission of Viamed Ltd.

Graph: SpO2 vs. SaO2
Viamed VM-2105 SpO2 Monitor vs. Co-oximetry (SaO2)
[image: image2.jpg]Sp0, %

100

95

90

85

80

75

70

65

o VM-2105 (integrated sensor)

Identity

------- Regression Line

70

75

80

85
Sa0, %

90

95

100




Bland-Altman Plot (difference: SpO2 – SaO2)
Viamed VM-2105 SpO2 Monitor: Difference vs. Co-oximetry (SaO2)
[image: image3.jpg][5p0, - 5a0,]

1o

o VM-2105 (integrated sensor)

Zero
+2 Sigma
-2 Sigma

— . — Bias

70

75

80

85
Sa0, %

90

95

100




Excluded Points(s):
(none)
VM-2105 (Grey) Clinical Study Report 

Date: 10/09


[image: image4.png]