VIAMED

SpO2 Probe Tester

VG900 Q.C. Inspection & Test Procedure

VM3/COP/55.02

	Date:10-Dec-03
	Revision Date:29-Dec-03
	Issue: 1



[image: image1.jpg]



Note: The facia panel of the VG-900 is very soft – it is easily scratched!  

Be extremely careful not to mark the facia panel during this procedure

Equipment Required:
1 x P856RA SpO2 Finger Probe – Nellcor





1 x P864RA SpO2 Finger Probe – Datascope





1 x P867RA SpO2 Finger Probe – Ohmeda





4 x “AA” Batteries





Isopropyl Alcohol

	Q.C. Inspection


1. Remove the units from the bubble wrap packaging material and arrange on a clean surface workbench.

2. Take each unit in turn and check for the following aesthetic conditions:

a) That there area no scratches or indentations on the outer case area.

b) That the Facia Panel is not scratched and that it is aligned correctly and evenly, within the recess.

c) That the 3 sockets are correctly aligned and orientated.

d) That the display window is not scratched and that the display is properly aligned and symmetrical within the window area.

e) That there are no loose bits of plastic showing in the window.

f) The rear panel stickers are not blurred or misaligned from the vertical.

3. If any units fail the above tests, then they must be returned to the supplier for rework (unless they are minor faults, and can be corrected on-site).

	Q.C. Testing


1.
Remove the battery panel and fit “4” x “AA” batteries, switch on and ensure the tester shows “4” Red LED’s on the “Open Circuit” line.

[image: image2.jpg]




[image: image3.jpg]



Picture No1



Picture No 2

[image: image4.jpg]




            [image: image5.jpg]Demonstration Model

'E% RESISTOR
LOW BATT. O







Picture No 3



 Picture No 4

2. Take each of the Finger Probes in turn and insert them into the appropriate socket on the panel.



    [image: image6.jpg]


   [image: image7.jpg]


   [image: image8.jpg]



3.
In each case, the display window should show “4” x Green LED’s.



[image: image9.jpg]Demonstration Model

OPEN
CIRCUIT $

DETECT

®rower





4. If any units fail i.e. “Open or Short-circuit indicators light up, then they must be forwarded to the workshop for internal repair and then re-test.

5. If the units pass, then switch off the unit and remove the batteries. Affix with a serial number from the stock sheet, and replace the battery cover.

6. Clean any marks from the case and re-wrap in its original bubble wrap packaging. Forward the completed batch to stock control.

4 x “AA”


Batteries





Battery


Panel





“4” Green


LED’s





Datascope





Nellcor





Ohmeda








“4” Red


LED’s





“4” Green


LED’s





Facia


Panel





Ohmeda





Nellcor





Datascope





On / Off


Switch








Page 3 of 3

