VM3COP50.13 Tom Thumb QA

| Required test equipment: |  |
| --- | --- |
| Snoop |  |
| Digital manometer |  |
| Tom Thumb QA test sheet |  |
| Precision valve isolation cover |  |

Quality Checks: Tom Thumb

Do not carry out QA checks on Tom Thumbs until at least 1 hour after testing – this allows Oxygen-Compatible thread-lock, applied during testing, to be totally dry.

- Check the Tom Thumb for damage i.e. scratches to plastics, poor chroming, etc.

- Set the adjustable valve to minimum (fully counter clockwise).

- Connect the inlet to the oxygen supply at a pressure of 4 bar.

- Connect the digital manometer (on 0-100.0 mbar range) with T piece adapter to the Tom Thumb outlet.

- Set the flow of oxygen to 10 lpm.

- Check all mechanical connections between parts for leaks; use Snoop fluid and check for bubbling.

- Check for minimum pressure out: Cover the T piece adapter port. Check the TT pressure gauge is less than 8.0 cmH2O. Check that the digital manometer reading is less than 7.8mbar. Record the Tom Thumb pressure gauge value on the Tom Thumb Calibration / Test & QA Sheet.

- Test the adjustable valve: Set the adjustable valve control to maximum. Using the precision valve isolation cover, cover the holes in the precision valve. Using your thumb, cover the T piece adapter port. Ensure a reading of 42.1±1 mbar on the digital manometer and record on the Tom Thumb Calibration / Test & QA Sheet.

- Check for TT pressure gauge sticking: Release and cover the T piece adapter port several times, ensuring that TT pressure gauge returns to 43±1 cmH2O. Remove the precision valve isolation cover.

- Test the precision valve: Using fingers and thumb, cover the holes in the adjustable valve and the T piece adapter port. Using the digital manometer, ensure that the Tom Thumb precision valve releases at 44.1±1mbar. The pressure gauge should read 45±1 cmH2O. Record the manometer value on the Tom Thumb Calibration / Test & QA Sheet.

- TT pressure gauge accuracy check: Cover the T piece adapter port and increase the flow of gas to achieve a displayed reading of 30 cmH2O on the TT pressure gauge. Ensure a reading of 29.4±1 mbar of the Precision Valve Reading on the digital manometer and record.

- Turn the flowmeter and adjustable valve to minimum. Disconnect from oxygen supply and disconnect the digital manometer.

- Clean the Tom Thumb with isopropyl alcohol.

- Ensure the appropriate labels have been fitted and the Instructions for Use card is attached. Complete stock sheets and file. Ensure the unit is booked onto the computerised stock system.