VM3COP44.01 - Envitec OxiQuant S QA Procedure

| Testing: |  |
| --- | --- |
| - Unbox the OxiQuant analyser (PN:0111275) and the Oxygen sensor (PN:0110660). |  |
| - Check the box contains all the items listed: 1 x OxiQuant analyser (PN:0111275) 1 x Oxygen sensor (PN:0110660) 1 x T piece adaptor 1 x Flow diverter 1 x Hose adaptor 1 x Operating instructions |  |
| - Attach the Oxygen sensor to analyser. Place the analyser over the sensor contacts and rotate the analyser clockwise, until locked in place. |  |
| - Activate the analyser by pressing and holding the power button for 3 seconds. - The display will briefly activate all the LCD segments. Check these are all working. - The unit will then briefly display the version number, before showing an Oxygen reading. |  |
| - The initial reading will start at 20.9%. This might fluctuate slightly (±1%) in air. |  |
| - In turn, press the ▲ and ▼buttons. Observe that these raise and lower the displayed Oxygen value. - Return the displayed value to 20.9%. |  |
| - The unit can be switched off by pressing the power button for 3 seconds. - Observe that the analyser automatically powers off after 3 minutes. Nothing is shown on the display when the analyser is switched off. |  |
| Packing and labelling: |  |
| - Disconnect the Oxygen sensor from the analyser by turning the analyser counter clockwise to the sensor. |  |
| - Check the serial number of the analyser, which is located on the underside of the unit. |  |
| - Confirm that the serial number of the unit matches that of the one found on the label on the outer packaging. |  |
| - Check the serial number of the Oxygen sensor, which can be found on the white label near the sensor base. |  |
| - Affix the corresponding serial number labels, for both the analyser and sensor, to the front face of the outer packaging. |  |
| - Record the QA results, for both the analyser and the sensor individually, in the Intrastats system. |  |