	      [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: 01535 634542               Fax: 01535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



Viamed 3700 Series Temperature Probes

R-T Specifications

	Part number
	Model
	Connector
	Type
	Applications

	0210400
	37204
	6.35mm (1/4") 
Jack (mono)
	General purpose 

Rectal/Oesophageal
	Paediatric/Neonatal, Veterinary, Pharmacology

	0210401
	37104
	6.35mm (1/4")  

Jack (mono)
	General purpose 

Rectal/Oesophageal
	Adult/Paediatric, Veterinary, Liquids

	0210403
	37804
	6.35mm (1/4")
Jack (mono)
	General purpose 

Skin/Surface
	Neonatal, Veterinary

	0210404
	37904
	6.35mm (1/4")  

Jack (mono)
	General purpose 

Skin/Surface
	Adult/Paediatric, Veterinary, Pharmacology

	0210410
	37204D
	3.5mm 

Jack  (mono)
	General purpose 

Rectal/Oesophageal
	Paediatric/Neonatal, Veterinary, Pharmacology

	0210411
	37104D
	3.5mm 

Jack (mono)
	General purpose 

Rectal/Oesophageal
	Adult/Paediatric, Veterinary, Liquids

	0210413
	37804D
	3.5mm 

Jack (mono)
	General purpose 

Skin/Surface
	Neonatal, Veterinary

	0210414
	37904D
	3.5mm 

Jack (mono)
	General purpose 

Skin/Surface
	Adult/Paediatric, Veterinary, Pharmacology

	Cable length: 3.00 m




	Resistance @ 25°C:     2252.00 Ω
Tolerance   @ 37°C:     +/- 0.1% 

Operating range:           -10°C to +80°C

Compatibility:                YSI 400 Series

     

	Temperature °C
	Resistance

Ω
	
	Temperature °C
	Resistance

Ω
	
	Temperature °C
	Resistance

Ω
	
	Temperature °C
	Resistance

Ω

	-10
	12453.00
	
	13
	3886.40
	
	36
	1411.10
	
	59
	580.68

	-9
	11796.00
	
	14
	3707.90
	
	37
	1354.40
	
	60
	560.09

	-8
	11177.00
	
	15
	3538.80
	
	38
	1300.20
	
	61
	540.34

	-7
	10595.00
	
	16
	3378.30
	
	39
	1248.50
	
	62
	521.38

	-6
	10046.00
	
	17
	3225.90
	
	40
	1199.20
	
	63
	503.18

	-5
	9529.20
	
	18
	3081.30
	
	41
	1152.00
	
	64
	485.70

	-4
	9011.70
	
	19
	2939.50
	
	42
	1107.00
	
	65
	468.94

	-3
	858.20
	
	20
	2813.60
	
	43
	1063.90
	
	66
	452.82

	-2
	8148.10
	
	21
	2689.60
	
	44
	1022.80
	
	67
	437.35

	-1
	7739.20
	
	22
	2571.80
	
	45
	983.42
	
	68
	422.48

	0
	7353.00
	
	23
	2459.80
	
	46
	945.81
	
	69
	408.20

	1
	6988.10
	
	24
	2353.30
	
	47
	909.83
	
	70
	394.47

	2
	6643.40
	
	25
	2252.00
	
	48
	875.40
	
	71
	381.26

	3
	6317.80
	
	26
	2155.60
	
	49
	842.47
	
	72
	368.57

	4
	6010.00
	
	27
	2063.90
	
	50
	810.94
	
	73
	356.36

	5
	5719.10
	
	28
	1976.60
	
	51
	780.75
	
	74
	344.62

	6
	5443.70
	
	29
	1893.40
	
	52
	751.85
	
	75
	333.32

	7
	5183.20
	
	30
	1814.20
	
	53
	724.16
	
	76
	322.46

	8
	4936.80
	
	31
	1738.80
	
	54
	697.65
	
	77
	311.99

	9
	4703.30
	
	32
	1666.90
	
	55
	672.24
	
	78
	301.92

	10
	4482.30
	
	33
	1598.30
	
	56
	647.88
	
	79
	292.23

	11
	4272.90
	
	34
	1532.90
	
	57
	624.53
	
	80
	282.89

	12
	4074.50
	
	35
	1470.60
	
	58
	602.15
	
	
	


[image: image1.png]