[image: image1.emf]

	COMPANY PROCEDURE
	CP NO:
	CPF29

	Title    
	REJECTION REPORT - CUSTOMER COMPLAINTS

	Issued By:
	BLE
	Dept:
	QA
	Clause No
	N/A

	Authorised By:
	MDE
	Issue
	001
	Revision
	004

	Date of Issue: APR 2017  -  ***DO NOT AMEND THIS TABLE***
	Page  1   of  1


TO BE COMPLETED BY GNW 

	RR No:
	RR171
	Date
	18-1-18

	Current Job No
	79038
	Cust No: 

(Office Only)
	

	Original Job No:
	78487
	Original Operator
	

	Cust Ref No:
	RTS110579
	GNW Order No:
	


	Source of Failure
	External (supplier)
	
	Carrier
	 

	Tick Box
	GNW Ltd
	
	Customer
	


	Details of Equipment
	1 x 63mm Dial Compound Gauge -5 to +60 cm/wtr

Centre back entry 1/8” NPT Connection
RTS1100579  

	Customer Reason for Complaint
	Stuck on 20 cm.Wg 

	Workshop Report
	Movement inside beyond repair, looks like gauge possibly dropped or over-pressured.

	Remedial Action taken
A remedial action is a change made to a nonconforming product or service to address the deficiency
	Replaced internal movement, fully tested up & calibrated.  Reading within specified accuracy (no errors)

	Corrective Action Report
Identification and elimination of the causes of a problem, thus preventing their recurrence
	 Handle gauge with care


Signed 





   Date:    



 

 GNW Instrumentation


    A division of Gauges North West Ltd

















