The intended use is to determine depth of Anesthesia by introducing electrical pulses into the muscle and looking for corresponding twitches / The Micostim DBS is not attached to the mains. It is a battery 9 V, 550mAh (PP3 Alkaline) operated instrument that develops stimuli of 200 µs pulse width up to ≥ 90 V p-p (100 mA) with 1 KΩ load, in order to assess the patient’s condition under anaesthesia. In the worst-case failures: The switch has independent on/off for battery and pulses. It needs physical action on behalf of the user to function. Continuous output the circuitry, assuming the switches are held on cannot sustain this output for more than a few seconds. Capacitor discharge is visible even during Train of four. The unit is battery powered with a limited capacity. There are no current specific standards for Supramaximal Peripheral Nerve Stimulators / It is used for diagnostic purposes and/or monitoring depth of Anaesthesia / It does so by by introducing electrical pulses into the muscle and looking for corresponding twitches	Inaccurate diagnosis could result in the inappropriate use of the equipment. The equipment should only be used by experienced trained personnel / The main risk is an non harmful but painful electric shock
It used by individuals with various skill levels from various cultural backgrounds?	The equipment should only be used by experienced trained personnel / Symbols are used
There is Contact with enclosure, controls & Leads by the patient or user whilst using the product. ECG pads not supplied with product – customers use their own. CE branded product. / Microstim Electrodes / The Microstim delivers an electric pulse to the skin. It is immaterial how the device is connected. It can be dumbbells, two wires taped to the skin or ECG electrodes. ECG electrodes were chosen because they were readily available in most clinical areas and are very familiar to staff. They are easy to use and offer consistent application. They were chosen in preference to Dumbbells because dumbbells only present a very small surface area and need to be used with extreme care. Intermittent contact can cause pain or distress to the patient. The Microstim can function adequately with two bare wires. Due to its design, it is reasonable to cope with small variations in patient skin impedance. It is designed to work on the universally accepted nominal around 1 KΩ. The use of ECG electrodes reduces poor contact problems. The function and efficacy of ECG electrodes is well documented as is they are universally used on ECG and patient monitoring equipment.	The contact is transient
Energy delivered to and/or extracted from the patient?	The equipment should only be used by experienced trained personnel / It can effect pacemakers ;Only one observation of this circa 1987 has been observed by a user Warning is given in the manual. Trained Anesthetists will be aware of this
Conclusion: The main risk is an non harmful but painful electric shock if the user inadvertently used a high setting or the leads had an intermittent fault. / Since 1986 there have been no problems and one occurrence with a pacemaker. This product is not unique and it has a range of competitors.
Probability / There are many thousands of Microstims daily operation. There have been no complaints and one observation that placement of either the Microstim or the electrodes could effect pacemakers .circa 1987. The only cause of failure is component. Probability is therefore based on the return/service figures. These are also very low considering the abuse they receive in Operating Theatres. The return figures show consistent reoccurring faults ( a single component) and sudden increaes of a particular fault which could indicate a batch problem of a change in the way the Microstim is being used. An overall cut off figure of 5% overall return/repair is used. / The overall probability outside of repair is remote.

Signed:	J.S.Lamb	Date 01/11/11