
Pt. [image: image1.png]


[image: image2.jpg]C€00s6



No 2390001


                                                                                Version v2.0


                                                                                                                                  [image: image2.jpg]                                                                                           


Instructions for use Disposable Oxygen Hoods/Tents

Part numbers 2310001 to 23100134
[image: image3.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




[image: image4.png]



Assembly and General Instructions:

1) Remove oxygen hood/tent from the packaging.

2) Unfold oxygen hood/tent and remove press-out discs from the 22mm tubing ports where present. In certain hoods/tents such as the Double-Tent Outer, the discs have been left in place. In other models they have already been removed.

3) Snap the back panel into the body section using all the snap connectors. Note: If the tent is a No-Bottom, NICU or Partial Bottom design, the snap connections on the bottom edge of the back panel are not used.

4) The three 22mm tubing ports in the back panel of the Outer Double-Tent are available for vacuum and aerosol/humidity tubing. Note: Tubing, connectors, and filters of the user’s choice will need to be added.

5) Assembled hoods/tents should be used according to the hospital’s protocols.

6) Bedding should be placed between the patient and the hood/tent bottom where it exists.

	[image: image6.jpg]



Warning.
	· Prior to placing the patient inside the hoods/tent, the hood/tent should be inspected to ensure that the press-out discs have been removed from the hood/tent and the hood/tent area. Ensure that all the snap connections are in working order.
· No smoking, no use of electrical appliances, no open flames should be allowed when oxygen is in use; only approved toys should be given to the patient
· When using all types of Viamed oxygen hoods, use at least 6 litres / minute of supplied gas.

· Carry out checks of the infants’ blood oxygen level to be certain of adequate CO2 washout.

· For babies of less than 1.5Kg, pre-warm the oxygen hood to approximately 35(C before use.

 Placing a cold oxygen hood over a small baby may decrease the body temperature as radiant-loss from the baby can contribute to the warming of the oxygen hood.


Viamed Ltd, 15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, UK.

Tel: +44 (0)1535 634542. Fax: +44 (0)1535 635582.

Email:  info@viamed.co.uk.  Website: www.viamed.co.uk.      

© Copyright 2007. Viamed Ltd.

[image: image5.jpg]


