Vandagraph

Replacement of the Analox O2 DII Oxygen sensor 9212

&

the Divex Mini O2
Using 

Vandagraph  R-33VAN

The following notes are a guide

[image: image1.jpg]


[image: image2.jpg]



[image: image3.jpg]



1. Refer to manufacturers instructions. 

2. Loosen the 4 screws located on the front cover.

3. Carefully lift the cover.

4. Disconnect the connector from the back of the sensor

5. Remove the sensor from the front cover.

6. Remove the new sensor from its pack and check it for leaks

7. Check the sensor has a rubber 0-ring fitted.

8. Introduce the sensor into the front cover 

9. [image: image4.jpg]


Make sure the metal prongs and sockets on the connectors line up and push in firmly.

10. Ensure the wires are not trapped under the sensor 

11. Replace the cover carefully taking care that the sensor locates properly and that you do not trap any wires

12. Do not over tighten the screws.

13. Leave exposed to dry outdoor air for at least two minutes NB (All galvanic sensors can take an hour to fully stabilize after they have first been exposed to air to achieve maximum accuracy. This may be visible as a rising output)

14. Calibrate to 20.9%

15. [image: image5.jpg]


Dispose of the old sensor according to local regulations for Lead and Potassium Hydroxide solution


Vandagraph Ltd.  15,Station Road, Crosshills , Keighley, West Yorkshire. BD20 7DT

+44 (0) 1535 634900   Fax: +44 (0) 1535 635582

www.vandagraph.co.uk      email  sales@vandagraph.co.uk
Analox O2DII sensor replacement 070305.doc
Page 1
07/03/2007

