Label part number 2590100

# Microstim DB3 Supramaximal Nerve Stimulator

Part number: 2510001

Serial number:  H000xxxx

#### Caution:

This is a non-standard version unit and is easily identified as it has green switches; the standard

version (part number 2510000) has red switches.

It is essential that the user reads the supplementary  instructions – Appendix I, page 20.

1 February  2010

Label part number 2590101

# Appendix I

Applies to DB3 Supramaximal Nerve Stimulators, part number 2510001

Functionality of stimulation is as follows:

T.O.F.   Latching switch – Continuous cyclic output, T.O.F. which repeats after a pause of

8 seconds - for a total period of 5 minutes.

D.B.S.  Momentary switch - Functionality same as with standard DB3

P.T.C.  Momentary switch - Functionality same  as with standard DB3

1 Hz .   Latching switch – Continuous output repeats for a total period of 5 minutes.

NOTE: With either of the revised T.O.F. and 1Hz functions the latching switches need to

be reset to the middle position otherwise another function cannot be selected.

After use - if either of the T.O.F and 1 Hz functions are left in the latched (On) position there is still a

small current drain, even after the five minute time-out period. This current consumption will drain the

battery unless the switches are returned to the middle (Off) position.

1st February 2010