[image: image1.png]


 

                   O2 Sensors
	COMPANY OPERATING PROCEDURES


Teledyne R21A (PN 0110121), 
O2 Sensor Test Procedure
VM3/COP/37.36
	Date:
  3-Feb-03
	Revision date: 12-Jul-18
	Issue:   3



Test & Equipment Required

Functional test required using R21 tester and R21 test adapter, (digital multimeter and adapter).

Output specification as in document: (VM3COP37.00, column (i)).

Method.

1. Open cell packaging and remove O2 sensor.

2. Check the O2 sensor for damage and signs of leaking electrolyte. 

3. Connect the O2 sensor to the R21 tester, select 200mV range (marked with arrows) and wait a few seconds for meter reading to stabilise.


4. Check for an output  Ensure that meter reading lies within the output specification range,
7mV - 14mV

5. If no, or very low output, or sensor is returned as faulty leave the O2 sensor exposed to air for 1 Hour Minimum.


6. If the O2 sensor fails the output specification test, it should be left unsealed and boxed, placed in a docket and put on the goods in desk with a short note explaining why it failed, ready to be booked in with a SRN

7. Disconnect the O2 sensor from the R21 tester and label with Viamed stickers if appropriate.

8. Reseal in serial numbered packet, integrity seal the edge of the packet with initialed and dated Viamed sticker and re-box. 

VM3COP37.36 R-21A Oxygen Sensor 0110121.doc _ID406.doc    12/07/2018         Page 1 of 1

[image: image1.png]