	[image: image3.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley - West Yorkshire BD20 7DT - United Kingdom
Tel: +44 (0)1535 634542         Fax: +44 (0)1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk




Parts Required for the Annual Service of Tom Thumb Infant Resuscitator

This information relates to the following models of Tom Thumb infant resuscitator:

· 0310030 - TT480 with low-pressure tapered gas inlet

· 0310032 - TT490-15 with integral flowmeter and 1 m oxygen hose

· 0310034 - TT490-15-3 with integral flowmeter and 3 m oxygen hose

To perform an annual service, the following parts are required: 


	Qty
	Part No
	Description

	1
	0330213
	O-ring, thick

	6
	0330214
	O-ring, thin

	The following may be required if the gauge is not sufficiently tight:

	1
	0330212
	Gauge seal (diamond copper washer)


Notes: 

Please see separate price list for current pricing. 

For any other variants of Tom Thumb, please contact Viamed. 

The gauge seal is only required if the gauge is loose or has been removed.

This seal is a diamond-profile copper washer that it designed to compress as the gauge is tightened, allowing the gauge to be rotated up to 1 full turn from initial contact with the washer so as to allow correct alignment. 


	[image: image1.jpg]



	CAUTION 

Do not re-use a gauge seal that has been previously compressed.


	[image: image2.jpg]



	CAUTION 

For variants of Tom Thumb that have an oxygen or air hose, Viamed recommends that these be replaced every 4 years.


Hoses are available in different lengths; please contact Viamed with the required gas type and hose length for a part number and price.

v1.2 – August 2018

[image: image3.png]